## HS-VOC
15-03379-N1 |  |  | HS-VOC
| Sample Mass (g) | Compound | Retention Time | Response | Area > 10 % | Conc (ug) | Conc (ug/g) | ID Level | CAS
Sample 1 | 4.78 | Ethylene Oxide | 1.56 | 5252650 | TRUE | 2.7713 | 0.57978 | TIC | 75-21-8
|  | Unknown | 2.69 | 284523 | TRUE | 0.15012 | 0.031405 | U | N/A
|  | Cyclohexane | 5.1 | 522608 | TRUE | 0.27573 | 0.057684 | MPC | 110-82-7
|  | Ethylene Benzene | 11.59 | 263870 | TRUE | 0.13922 | 0.029125 | MPC | 100-41-4
|  | p-xylene | 11.8 | 202187 | TRUE | 0.10668 | 0.022317 | TIC | 106-42-3
|  | Decane | 14.17 | 196653 | TRUE | 0.10376 | 0.021706 | MPC | 124-18-5
|  | Undecane | 16.11 | 329712 | TRUE | 0.17396 | 0.036393 | MPC | 1120-21-4
|  | Dodecane | 17.88 | 406980 | TRUE | 0.21473 | 0.044922 | MPC | 629-50-5
|  | Tridecane | 19.52 | 284356 | TRUE | 0.15003 | 0.031387 | MPC | 629-50-5
Sample 2 | 4.79 | Ethylene Oxide | 1.46 | 5382570 | TRUE | 2.8399 | 0.59288 | TIC | 75-21-8
|  | Unknown | 2.68 | 247449 | TRUE | 0.13056 | 0.027256 | U | N/A
|  | Cyclohexane | 5.09 | 568846 | TRUE | 0.30013 | 0.062657 | MPC | 110-82-7
|  | Ethylene Benzene | 11.6 | 272592 | TRUE | 0.14382 | 0.030025 | MPC | 100-41-4
|  | p-xylene | 11.8 | 236796 | TRUE | 0.12494 | 0.026082 | TIC | 106-42-3
|  | Decane | 14.17 | 122330 | FALSE | 0.064542 | 0.013474 | MPC | 124-18-5
|  | Undecane | 16.11 | 330961 | TRUE | 0.17462 | 0.036455 | MPC | 1120-21-4
|  | Unknown | 16.92 | 253812 | TRUE | 0.13391 | 0.027957 | U | N/A
|  | Dodecane | 17.88 | 436297 | TRUE | 0.23019 | 0.048057 | MPC | 629-50-5
|  | Naphthalene | 18.6 | 502096 | TRUE | 0.26491 | 0.055305 | MPC | 91-20-3
|  | 1,2,3-trichloro benzene | 19 | 298301 | TRUE | 0.15739 | 0.032857 | MPC | 87-61-6
|  | Tridecane | 19.52 | 274919 | TRUE | 0.14505 | 0.030282 | TIC | 629-50-5
Sample 3 | 4.79 | Ethylene Oxide | 1.46 | 5526490 | TRUE | 2.9158 | 0.60873 | TIC | 75-21-8
|  | Unknown | 2.69 | 251825 | TRUE | 0.13286 | 0.027738 | U | N/A
|  | Cyclohexane | 5.09 | 612789 | TRUE | 0.32331 | 0.067497 | MPC | 110-82-7
|  | Ethylene Benzene | 11.6 | 283055 | TRUE | 0.14934 | 0.031178 | MPC | 100-41-4
|  | p-xylene | 11.8 | 171869 | FALSE | 0.090679 | 0.018931 | TIC | 106-42-3
|  | Decane | 14.17 | 191250 | TRUE | 0.1009 | 0.021066 | MPC | 124-18-5
|  | Undecane | 16.11 | 403632 | TRUE | 0.21296 | 0.044459 | MPC | 1120-21-4
|  | Dodecane | 17.88 | 501143 | TRUE | 0.26441 | 0.0552 | MPC | 629-50-5
|  | Tridecane | 19.52 | 342529 | TRUE | 0.18072 | 0.037729 | TIC | 629-50-5
|  |  | Area Internal | 1895353 | Run ID: | A0074
|  |  | 10% area internal | 189535.3
|  |  | Conc (ug) | 1
Conc (ug) = Conc Standard (ug) * Area Sample/Area Standard
Conc (ug/g) = Conc (ug)/Sample Mass (g)
Highlighted results reported.
MPC - >80%
TIC - <80%
Note on B14: sdoherty: Update formula in Column G

## VOC
15-03379-N1 |  | VOC
[7 empty] | Control | Reportable
|  |  |  | Instrument | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/L) | Factor | Result (ug/L) | Result (ug/L) | (R/NR) | ID | CAS #
Drug Solution | Control | Isopropyl Alcohol | 3.135 | 1191.56 | 5000 | 5957800 | 0 | - | IC | 67-63-0
| Syringe S1 | Isopropyl Alcohol | 3.124 | 1048.28 | 5000 | 5241400 | -716400 | NR | IC | 67-63-1
| Syringe S2 | Isopropyl Alcohol | 3.124 | 1097.4 | 5000 | 5487000 | -470800 | NR | IC | 67-63-2
| Syringe S3 | Isopropyl Alcohol | 3.135 | 1242.34 | 5000 | 6211700 | 253900 | NR | IC | 67-63-3
| Accu-Chek S1 | Isopropyl Alcohol | 3.135 | 969.73 | 5000 | 4848650 | -1109150 | NR | IC | 67-63-1
| Accu-Chek S2 | Isopropyl Alcohol | 3.135 | 1029.46 | 5000 | 5147300 | -810500 | NR | IC | 67-63-2
| Accu-Chek S3 | Isopropyl Alcohol | 3.135 | 1063.05 | 5000 | 5315250 | -642550 | NR | IC | 67-63-3
| Neria S1 | Isopropyl Alcohol | 3.135 | 1056.61 | 5000 | 5283050 | -674750 | NR | IC | 67-63-1
| Neria S2 | Isopropyl Alcohol | 3.135 | 1072.59 | 5000 | 5362950 | -594850 | NR | IC | 67-63-2
|  | Cyclopentasiloxane, decamethyl- | 18.3 | 8.62 | 5000 | 43100 | 43100 | R | TIC | 541-02-6
| Neria S3 | Isopropyl Alcohol | 3.124 | 1088.91 | 5000 | 5444550 | -513250 | NR | IC | 67-63-3
15-03379-N1 |  | VOC
[7 empty] | Control | Reportable
|  |  |  | Instrument | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/L) | Factor | Result (ug/L) | Result (ug/L) | (R/NR) | ID | CAS #
Acidic Water | Control | N/A | N/A | N/A | N/A | N/A | N/A | - | N/A | N/A
| Syringe S1 | Acetaldehyde | 1.32 | 7.91 | 1 | 7.91 | 7.91 | R | TIC | 75-07-0
|  | Acetone | 2.24 | 493.08 | 1 | 493.08 | 493.08 | NR | IC | 67-64-1
|  | tert-Butyl alcohol | 2.87 | 5946.39 | 1 | 5946.39 | 5946.39 | NR | IC | 75-65-0
|  | methyl tert-butyl ether | 2.98 | 36.04 | 1 | 36.04 | 36.04 | R | IC | 1634-04-4
|  | 1-Pentene, 2-methyl- | 3.17 | 5.6 | 1 | 5.6 | 5.6 | NR | MPC | 763-29-1
|  | 2-Butenal, (E)- | 3.45 | 7.22 | 1 | 7.22 | 7.22 | NR | MPC | 123-73-9
|  | 2-Butanone | 4.51 | 11.3 | 1 | 11.3 | 11.3 | R | IC | 78-93-3
|  | 5-Hepten-2-one | 18.8 | 3.91 | 1 | 3.91 | 3.91 | R | TIC | 1071-94-9
|  | Unknown | 19.03 | 12.55 | 1 | 12.55 | 12.55 | R | U | N/A
| Syringe S2 | Acetaldehyde | 1.32 | 6.19 | 1 | 6.19 | 6.19 | R | TIC | 75-07-0
|  | Acetone | 2.24 | 474.31 | 1 | 474.31 | 474.31 | NR | IC | 67-64-1
|  | tert-Butyl alcohol | 2.86 | 4944.29 | 1 | 4944.29 | 4944.29 | NR | IC | 75-65-0
|  | methyl tert-butyl ether | 2.98 | 28.86 | 1 | 28.86 | 28.86 | R | IC | 1634-04-4
|  | 1-Pentene, 2-methyl- | 3.17 | 5.1 | 1 | 5.1 | 5.1 | NR | TIC | 763-29-1
|  | 2-Butenal, (E)- | 3.45 | 6.28 | 1 | 6.28 | 6.28 | NR | MPC | 123-73-9
|  | 2-Butanone | 4.49 | 9.24 | 1 | 9.24 | 9.24 | R | IC | 78-93-3
|  | 5-Hepten-2-one | 18.8 | 6.96 | 1 | 6.96 | 6.96 | R | TIC | 1071-94-9
|  | Unknown | 19.03 | 50.7 | 1 | 50.7 | 50.7 | R | U | N/A
| Syringe S3 | Acetaldehyde | 1.32 | 6.76 | 1 | 6.76 | 6.76 | R | TIC | 75-07-0
|  | Acetone | 2.24 | 459.11 | 1 | 459.11 | 459.11 | NR | IC | 67-64-1
|  | tert-Butyl alcohol | 2.86 | 4700.21 | 1 | 4700.21 | 4700.21 | NR | IC | 75-65-0
|  | methyl tert-butyl ether | 2.96 | 22.52 | 1 | 22.52 | 22.52 | R | IC | 1634-04-4
|  | 1-Pentene, 2-methyl- | 3.16 | 4.54 | 1 | 4.54 | 4.54 | NR | TIC | 763-29-1
|  | 2-Butenal, (E)- | 3.45 | 5.86 | 1 | 5.86 | 5.86 | NR | MPC | 123-73-9
|  | 2-Butanone | 4.51 | 8.12 | 1 | 8.12 | 8.12 | R | IC | 78-93-3
|  | 5-Hepten-2-one | 18.8 | 6.97 | 1 | 6.97 | 6.97 | R | TIC | 1071-94-9
|  | Unknown | 19.03 | 50.64 | 1 | 50.64 | 50.64 | R | U | N/A
| Accu-Chek S1 | Chloromethane | 1.16 | 41.19 | 1 | 41.19 | 41.19 | NR | IC | 74-87-3
|  | Acetone | 2.24 | 128.1 | 1 | 128.1 | 128.1 | NR | IC | 67-64-1
|  | Isopropyl Alcohol | 2.42 | 491.48 | 1 | 491.48 | 491.48 | NR | IC | 67-63-0
|  | Tert-Butyl alcohol | 2.86 | 993.1 | 1 | 993.1 | 993.1 | NR | IC | 75-65-0
|  | 2-Propenoic acid, methyl ester | 4.67 | 20.16 | 1 | 20.16 | 20.16 | NR | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 95.6 | 1 | 95.6 | 95.6 | NR | MPC | 80-62-6
| Accu-Chek S2 | Chloromethane | 1.16 | 48.75 | 1 | 48.75 | 48.75 | R | IC | 74-87-3
|  | Acetone | 2.24 | 37.43 | 1 | 37.43 | 37.43 | R | IC | 67-64-1
|  | tert-Butyl alcohol | 2.86 | 220.19 | 1 | 220.19 | 220.19 | R | IC | 75-65-0
|  | 2-Propenoic acid, methyl ester | 4.67 | 21.66 | 1 | 21.66 | 21.66 | R | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 84.04 | 1 | 84.04 | 84.04 | R | MPC | 80-62-6
| Accu-Chek S3 | Chloromethane | 1.16 | 28.03 | 1 | 28.03 | 28.03 | R | IC | 74-87-3
|  | Acetone | 2.23 | 28.13 | 1 | 28.13 | 28.13 | R | IC | 67-64-1
|  | tert-Butyl alcohol | 2.86 | 99.86 | 1 | 99.86 | 99.86 | R | IC | 75-65-0
|  | 2-Propenoic acid, methyl ester | 4.66 | 20.92 | 1 | 20.92 | 20.92 | R | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 83.91 | 1 | 83.91 | 83.91 | R | MPC | 80-62-6
| Neria S1 | Chloromethane | 1.16 | 16.12 | 1 | 16.12 | 16.12 | R | IC | 74-87-3
|  | Acetaldehyde | 1.32 | 1.608 | 1 | 1.608 | 1.608 | R | TIC | 75-07-0
|  | Propanal | 2.15 | 1.89 | 1 | 1.89 | 1.89 | R | TIC | 123-38-6
|  | Acetone | 2.24 | 21.8 | 1 | 21.8 | 21.8 | R | IC | 67-64-1
|  | 2-Propenoic acid, methyl ester | 4.66 | 25.94 | 1 | 25.94 | 25.94 | R | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 68.38 | 1 | 68.38 | 68.38 | R | TIC | 80-62-6
Neria S2
|  | Chloromethane | 1.16 | 12.96 | 1 | 12.96 | 12.96 | R | IC | 74-87-3
|  | Acetaldehyde | 1.32 | 6.83 | 1 | 6.83 | 6.83 | R | MPC | 75-07-0
|  | Propanal | 2.15 | 9.84 | 1 | 9.84 | 9.84 | R | TIC | 123-38-6
|  | Acetone | 2.24 | 18.94 | 1 | 18.94 | 18.94 | R | IC | 67-64-1
|  | 2-Propenoic acid, methyl ester | 4.64 | 26.34 | 1 | 26.34 | 26.34 | R | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 61.65 | 1 | 61.65 | 61.65 | R | TIC | 80-62-6
| Neria S3 | Chloromethane | 1.16 | 22.4 | 1 | 22.4 | 22.4 | R | IC | 74-87-3
|  | Acetaldehyde | 1.32 | 10.15 | 1 | 10.15 | 10.15 | R | MPC | 75-07-0
|  | Propananl | 2.15 | 13.3 | 1 | 13.3 | 13.3 | R | TIC | 123-38-6
|  | Acetone | 2.24 | 29.55 | 1 | 29.55 | 29.55 | R | IC | 67-64-1
|  | 2-Propenoic acid, methyl ester | 4.66 | 34.1 | 1 | 34.1 | 34.1 | R | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 84.35 | 1 | 84.35 | 84.35 | R | TIC | 80-62-6
| Accu-Chek S1 Re-run | Chloromethane | 1.16 | 39.57 | 1 | 39.57 | 39.57 | R | IC | 74-87-3
|  | Acetone | 2.24 | 35 | 1 | 35 | 35 | R | IC | 67-64-1
|  | Isopropyl Alcohol | 2.43 | 386.77 | 1 | 386.77 | 386.77 | R | IC | 67-63-0
|  | tert-Butyl alcohol | 2.86 | 44.48 | 1 | 44.48 | 44.48 | R | IC | 75-65-0
|  | 2-Propenoic acid, methyl ester | 4.67 | 18.07 | 1 | 18.07 | 18.07 | R | TIC | 96-33-3
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.98 | 82.59 | 1 | 82.59 | 82.59 | R | MPC | 80-62-6
| Syringe S1 DL5 | Acetone | 2.24 | 90.06 | 5 | 450.3 | 450.3 | R | IC | 67-64-1
|  | tert-Butyl alcohol | 2.86 | 947.79 | 5 | 4738.95 | 4738.95 | R | IC | 67-63-0
|  | methyl tert-butyl ether | 2.97 | 6.1 | 5 | 30.5 | 30.5 | NR | IC | 1634-04-4
| Syringe S2 DL5 | Acetone | 2.24 | 85.3 | 5 | 426.5 | 426.5 | R | IC | 67-64-1
|  | tert-Butyl alcohol | 2.87 | 815.64 | 5 | 4078.2 | 4078.2 | R | IC | 67-63-0
|  | methyl tert-butyl ether | 2.98 | 5.7 | 5 | 28.5 | 28.5 | NR | IC | 1634-04-4
| Syringe S3 DL5 | Acetone | 2.24 | 75.7 | 5 | 378.5 | 378.5 | R | IC | 67-64-1
|  | tert-Butyl alcohol | 2.87 | 752.29 | 5 | 3761.45 | 3761.45 | R | IC | 67-63-0
|  | methyl tert-butyl ether | 2.98 | 4 | 5 | 20 | 20 | NR | IC | 1634-04-4
15-03379-N1 |  | VOC
[7 empty] | Control | Reportable
|  |  |  | Instrument | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/L) | Factor | Result (ug/L) | Result (ug/L) | (R/NR) | ID | CAS #
IPA | Control DL100 | Acetone | 2.25 | 20.51 | 100 | 2051 | 2051 | - | IC | 67-64-1
| IPA Matrix Blank DL100 | Acetone | 2.27 | 44.68 | 100 | 4468 | 4468 | - | IC | 67-64-1
| Accu-Chek S1 DL100 | Cloromethane | 1.16 | 4.03 | 100 | 403 | 403 | NR | IC | 74-87-3
|  | Acetone | 2.25 | 39.34 | 100 | 3934 | 1883 |  | IC | 67-64-1
|  | ethyl tert-butyl ether | 4.27 | 5.44 | 100 | 544 | 544 | NR | IC | 1634-04-4
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 8.01 | 382.18 | 100 | 38218 | 38218 | R | MPC | 80-62-6
|  | toluene | 9.4 | 2.6 | 100 | 260 | 260 | NR | IC | 108-88-3
|  | ethylbenzene | 11.89 | 2.47 | 100 | 247 | 247 | NR | IC | 100-41-4
|  | Styrene | 12.74 | 56.62 | 100 | 5662 | 5662 | R | IC | 100-42-5
|  | Heptane, 2,2,4,6,6-pentamethyl- | 14.35 | 25.12 | 100 | 2512 | 2512 | R | MPC | 13475-82-6
|  | Benzene, 1,3-bis(1-methylethyl)- | 17.74 | 31.39 | 100 | 3139 | 3139 | R | MPC | 99-62-7
| Accu-Chek S2 DL100 | Cloromethane | 1.16 | 3.92 | 100 | 392 | 392 | NR | IC | 74-87-3
|  | Acetone | 2.25 | 56.28 | 100 | 5628 | 3577 |  | IC | 67-64-1
|  | ethyl tert-butyl ether | 4.27 | 5.26 | 100 | 526 | 526 | NR | IC | 1634-04-4
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 8.01 | 392.91 | 100 | 39291 | 39291 | R | MPC | 80-62-6
|  | toluene | 9.4 | 2.5 | 100 | 250 | 250 | NR | IC | 108-88-3
|  | ethylbenzene | 11.89 | 2.77 | 100 | 277 | 277 | NR | IC | 100-41-4
|  | Styrene | 12.74 | 60.17 | 100 | 6017 | 6017 | R | IC | 100-42-5
|  | Heptane, 2,2,4,6,6-pentamethyl- | 14.36 | 25.75 | 100 | 2575 | 2575 | R | TIC | 13475-82-6
|  | Benzene, 1,3-bis(1-methylethyl)- | 17.74 | 32.66 | 100 | 3266 | 3266 | R | MPC | 99-62-7
|  | Benzene, 1,4-bis(1-methylethyl)- | 18.1 | 19.46 | 100 | 1946 | 1946 | R | MPC | 100-18-5
| Accu-Chek S3 DL100 | Cloromethane | 1.16 | 4.05 | 100 | 405 | 405 | NR | IC | 74-87-3
|  | Acetone | 2.27 | 60.19 | 100 | 6019 | 3968 |  | IC | 67-64-1
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 8 | 389.55 | 100 | 38955 | 38955 | R | MPC | 80-62-6
|  | toluene | 9.4 | 2.97 | 100 | 297 | 297 | NR | IC | 108-88-3
|  | ethylbenzene | 11.89 | 2.83 | 100 | 283 | 283 | NR | IC | 100-41-4
|  | Styrene | 12.74 | 64.1 | 100 | 6410 | 6410 | R | IC | 100-42-5
|  | Heptane, 2,2,4,6,6-pentamethyl- | 14.35 | 32.21 | 100 | 3221 | 3221 | R | MPC | 13475-82-6
|  | Benzene, 1,3-bis(1-methylethyl)- | 17.74 | 35.82 | 100 | 3582 | 3582 | R | MPC | 99-62-7
| Neria S1 DL100 | Cloromethane | 1.16 | 1.29 | 100 | 129 | 129 | NR | IC | 74-87-3
|  | Acetone | 2.27 | 55.86 | 100 | 5586 | 3535 |  | IC | 67-64-1
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 8 | 328.37 | 100 | 32837 | 32837 | R | MPC | 80-62-6
|  | toluene | 9.4 | 1.88 | 100 | 188 | 188 | NR | IC | 108-88-3
|  | ethylbenzene | 11.88 | 2.25 | 100 | 225 | 225 | NR | IC | 100-41-4
|  | Styrene | 12.74 | 47.05 | 100 | 4705 | 4705 | R | IC | 100-42-5
|  | Heptane, 2,2,4,6,6-pentamethyl- | 14.35 | 24.9 | 100 | 2490 | 2490 | R | MPC | 13475-82-6
|  | Benzene, 1,3-bis(1-methylethyl)- | 17.74 | 33.87 | 100 | 3387 | 3387 | R | MPC | 99-62-7
|  | Benzene, 1,4-bis(1-methylethyl)- | 18.1 | 18.97 | 100 | 1897 | 1897 | R | MPC | 100-18-5
| Neria S2 DL100 | Cloromethane | 1.16 | 1.35 | 100 | 135 | 135 | NR | IC | 74-87-3
|  | Acetone | 2.25 | 41.02 | 100 | 4102 | 2051 |  | IC | 67-64-1
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.99 | 326.01 | 100 | 32601 | 32601 | R | MPC | 80-62-6
|  | toluene | 9.4 | 2.23 | 100 | 223 | 223 | NR | IC | 108-88-3
|  | ethylbenzene | 11.88 | 2.59 | 100 | 259 | 259 | NR | IC | 100-41-4
|  | Styrene | 12.74 | 55.42 | 100 | 5542 | 5542 | R | IC | 100-42-5
|  | Heptane, 2,2,4,6,6-pentamethyl- | 14.35 | 29.94 | 100 | 2994 | 2994 | R | MPC | 13475-82-6
|  | Benzene, 1,3-bis(1-methylethyl)- | 17.73 | 38.62 | 100 | 3862 | 3862 | R | MPC | 99-62-7
|  | Benzene, 1,4-bis(1-methylethyl)- | 18.1 | 21.39 | 100 | 2139 | 2139 | R | MPC | 100-18-5
| Neria S3 DL100 | Cloromethane | 1.16 | 1.42 | 100 | 142 | 142 | NR | IC | 74-87-3
|  | Acetone | 2.26 | 36.32 | 100 | 3632 | 1581 |  | IC | 67-64-1
|  | 2-Propenoic acid, 2-methyl-, methyl ester | 7.99 | 319.83 | 100 | 31983 | 31983 | R | MPC | 80-62-6
|  | toluene | 9.39 | 1.91 | 100 | 191 | 191 | NR | IC | 108-88-3
|  | ethylbenzene | 11.88 | 2.27 | 100 | 227 | 227 | NR | IC | 100-41-4
|  | Styrene | 12.74 | 46.57 | 100 | 4657 | 4657 | R | IC | 100-42-5
|  | Heptane, 2,2,4,6,6-pentamethyl- | 14.35 | 26.08 | 100 | 2608 | 2608 | R | MPC | 13475-82-6
|  | Unkown | 15.42 | 16.6 | 100 | 1660 | 1660 | R | U | N/A
|  | Benzene, 1,3-bis(1-methylethyl)- | 17.73 | 31.13 | 100 | 3113 | 3113 | R | MPC | 99-62-7
|  | Benzene, 1,4-bis(1-methylethyl)- | 18.1 | 16.94 | 100 | 1694 | 1694 | R | MPC | 100-18-5
|  | Butylated Hydroxytoluene | 20.23 | 27.95 | 100 | 2795 | 2795 | R | MPC | 128-37-0
| Syringe S1 DL100 | Acetone | 2.27 | 191.22 | 100 | 19122 | 17071 |  | IC | 67-64-1
|  | 2,4-Dimethyl-1-heptene | 10.95 | 27.66 | 100 | 2766 | 2766 | R | MPC | 19549-87-2
|  | 2,4,6-trimethyl-decane | 15.59 | 17.27 | 100 | 1727 | 1727 | R | TIC | 62108-27-4
|  | Butylated Hydroxytoluene | 20.23 | 19.26 | 100 | 1926 | 1926 | R | MPC | 128-37-0
|  | Unknown | 20.75 | 75.64 | 100 | 7564 | 7564 | R | U | N/A
|  | Unknown | 20.99 | 64.36 | 100 | 6436 | 6436 | R | U | N/A
|  | Unknown | 21.07 | 26.21 | 100 | 2621 | 2621 | R | U | N/A
|  | Himachala-2,4-diene | 21.21 | 20.55 | 100 | 2055 | 2055 | R | TIC | 60909-27-5
| Syringe S2 DL100 | Acetone | 2.27 | 190.53 | 100 | 19053 | 17002 |  | IC | 67-64-1
|  | 2,4-Dimethyl-1-heptene | 10.95 | 29.69 | 100 | 2969 | 2969 | R | MPC | 19549-87-2
|  | 2,4,6-trimethyl-decane | 15.59 | 17.63 | 100 | 1763 | 1763 | R | TIC | 62108-27-4
|  | Unknown | 19.03 | 24.62 | 100 | 2462 | 2462 | R | U | N/A
|  | Unknown | 20.75 | 78.52 | 100 | 7852 | 7852 | R | U | N/A
|  | Unknown | 20.98 | 59.91 | 100 | 5991 | 5991 | R | U | N/A
|  | Unknown | 21.07 | 23.27 | 100 | 2327 | 2327 | R | U | N/A
|  | Himachala-2,4-diene | 21.21 | 18.42 | 100 | 1842 | 1842 | R | MPC | 60909-27-5
| Syringe S3 DL100 | Acetone | 2.25 | 157.46 | 250 | 39365 | 37314 |  | IC | 67-64-1
|  | 2,4-Dimethyl-1-heptene | 10.95 | 27.4 | 250 | 6850 | 6850 | R | MPC | 19549-87-2
|  | 2,4,6-trimethyl-decane | 15.59 | 14.06 | 250 | 3515 | 3515 | R | TIC | 62108-27-4
|  | Unknown | 19.03 | 15.83 | 250 | 3957.5 | 3957.5 | R | U | N/A
|  | Unknown | 20.75 | 82.07 | 250 | 20517.5 | 20517.5 | R | U | N/A
|  | Unknown | 20.99 | 54.08 | 250 | 13520 | 13520 | R | U | N/A
|  | Unknown | 21.07 | 19.3 | 250 | 4825 | 4825 | R | U | N/A
|  | Himachala-2,4-diene | 21.21 | 16.31 | 250 | 4077.5 | 4077.5 | R | TIC | 60909-27-5
Corrected Result (ug/L) = Instrument Result (ug/L) * Dilution Factor - Control Result (ug/L)
MPC - >80%
TIC - <80%

## SVOC
15-03379-N1 |  | SVOC
[8 empty] | Blank | Reportable
Acidified Water |  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
| mb1511020OE | 4-hydroxy-4-methyl-2-Pentanone | 7.92 | 19.34 | 25 | 1 | 0.7736 | 0.7736 | NR | TIC | 123-42-2
| Control | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 19.82 | 25 | 1 | 0.7928 | 0.0192 | NR | TIC | 123-42-2
| Syringe S1 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 19.34 | 25 | 1 | 0.7736 | 0 | NR | TIC | 123-42-2
|  | 2,5-dimethyl-2,5-Hexanediol | 12.94 | 37.66 | 25 | 1 | 1.5064 | 1.5064 | R | TIC | 110-03-2
| Syringe S2 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 17.74 | 25 | 1 | 0.7096 | -0.064 | NR | TIC | 123-42-2
|  | 2,5-dimethyl-2,5-Hexanediol | 12.94 | 40.02 | 25 | 1 | 1.6008 | 1.6008 | R | TIC | 110-03-2
| Syringe S3 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 15.88 | 25 | 1 | 0.6352 | -0.1384 | NR | TIC | 123-42-2
|  | 2,5-dimethyl-2,5-Hexanediol | 12.94 | 37.04 | 25 | 1 | 1.4816 | 1.4816 | R | TIC | 110-03-2
| Accu-Chek S1 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 10.02 | 25 | 2 | 0.8016 | 0.028 | NR | TIC | 123-42-2
|  | 1,6-dioxacyclododecane-7,12-dione | 22.07 | 193.23 | 25 | 2 | 15.4584 | 15.4584 | R | MPC | 777-95-7
| Accu-Chek S2 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 9.71 | 25 | 2 | 0.7768 | 0.0032 | NR | TIC | 123-42-2
|  | 1,6-dioxacyclododecane-7,12-dione | 22.07 | 192.54 | 25 | 2 | 15.4032 | 15.4032 | R | MPC | 777-95-7
| Accu-Chek S3 | 4-hydroxy-4-methyl-2-Pentanone | 7.88 | 9.32 | 25 | 2 | 0.7456 | -0.028 | NR | TIC | 123-42-2
|  | 1,6-dioxacyclododecane-7,12-dione | 22.07 | 181.93 | 25 | 2 | 14.5544 | 14.5544 | R | MPC | 777-95-7
| Neria S1 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 9.67 | 25 | 2 | 0.7736 | 0 | NR | TIC | 123-42-2
|  | 1,6-dioxacyclododecane-7,12-dione | 22.08 | 222.42 | 25 | 2 | 17.7936 | 17.7936 | R | MPC | 777-95-7
| Neria S2 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 9.9 | 25 | 2 | 0.792 | 0.0184 | NR | TIC | 123-42-2
|  | 1,6-dioxacyclododecane-7,12-dione | 22.07 | 219.23 | 25 | 2 | 17.5384 | 17.5384 | R | MPC | 777-95-7
| Neria S2 | 4-hydroxy-4-methyl-2-Pentanone | 7.91 | 8.11 | 25 | 2 | 0.6488 | -0.1248 | NR | TIC | 123-42-2
|  | 1,6-dioxacyclododecane-7,12-dione | 22.07 | 213.48 | 25 | 2 | 17.0784 | 17.0784 | R | MPC | 777-95-7
15-03379-N1 |  | SVOC
[8 empty] | Blank | Reportable
IPA |  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
| mb1511020OE | 4-hydroxy-4-methyl-2-Pentanone | 7.98 | 20.65 | 25 | 1 | 0.826 | 0.826 | NR | TIC | 123-42-2
| Control | 4-hydroxy-4-methyl-2-Pentanone | 7.98 | 22.57 | 25 | 1 | 0.9028 | 0.9028 | NR | TIC | 123-42-2
.
| Accu-Chek d10 S1 | 1,4-Butanediol | 10.52 | 10.51 | 25 | 10 | 4.204 | 4.204 | R | MPC | 110-63-4
|  | 1-[4-(1-methylethyl)phenyl]-Ethanone | 18.08 | 10.12 | 25 | 10 | 4.048 | 4.048 | R | MPC | 645-13-6
|  | 1,6-Dioxacyclododecane-7,12-dione | 22.05 | 80.22 | 25 | 10 | 32.088 | 32.088 | R | MPC | 777-95-7
|  | Dodecyl acrylate | 23.85 | 31.53 | 25 | 10 | 12.612 | 12.612 | R | TIC | 2156-97-0
|  | (1-hydroxycyclohexyl)phenyl-Methanone | 24.19 | 20.24 | 25 | 10 | 8.096 | 8.096 | R | MPC | 947-19-3
|  | Unknown | 24.3 | 15.43 | 25 | 10 | 6.172 | 6.172 | R | U | -
|  | Hexadecanoic acid, methyl ester | 26.48 | 67.38 | 25 | 10 | 26.952 | 26.952 | R | MPC | 112-39-0
|  | Styrene-acrylonitrile Trimer | 26.93 | 10.65 | 25 | 10 | 4.26 | 4.26 | R | TIC | 57964-39-3
|  | Unknown | 27.13 | 21.79 | 25 | 10 | 8.716 | 8.716 | R | U | -
|  | Oleanitrile | 28.02 | 10.77 | 25 | 10 | 4.308 | 4.308 | R | MPC | 112-91-4
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.1 | 78.07 | 25 | 10 | 31.228 | 31.228 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.15 | 141.86 | 25 | 10 | 56.744 | 56.744 | R | U | -
|  | Octadecanoic Acid methyl ester | 28.3 | 178.19 | 25 | 10 | 71.276 | 71.276 | R | MPC | 112-61-8
|  | Oleic Acid | 28.44 | 22.81 | 25 | 10 | 9.124 | 9.124 | R | TIC | 112-80-1
|  | Octadecanoic Acid | 28.58 | 25.74 | 25 | 10 | 10.296 | 10.296 | R | MPC | 57-11-4
|  | Unknown | 29.56 | 34.84 | 25 | 10 | 13.936 | 13.936 | R | U | -
|  | Unknown | 29.69 | 19.03 | 25 | 10 | 7.612 | 7.612 | R | U | -
| Accu-Chek d10 S2 | 1,4-Butanediol | 10.43 | 11.51 | 25 | 10 | 4.604 | 4.604 | R | MPC | 110-63-4
|  | 1-[4-(1-methylethyl)phenyl]-Ethanone | 18.08 | 10.15 | 25 | 10 | 4.06 | 4.06 | R | MPC | 645-13-6
|  | 1,6-Dioxacyclododecane-7,12-dione | 22.05 | 80 | 25 | 10 | 32 | 32 | R | MPC | 777-95-7
|  | Dodecyl acrylate | 23.85 | 31.77 | 25 | 10 | 12.708 | 12.708 | R | TIC | 2156-97-0
|  | (1-hydroxycyclohexyl)phenyl-Methanone | 24.19 | 20.97 | 25 | 10 | 8.388 | 8.388 | R | MPC | 947-19-3
|  | Unknown | 24.3 | 14.81 | 25 | 10 | 5.924 | 5.924 | R | U | -
|  | Hexadecanoic acid, methyl ester | 26.48 | 66.6 | 25 | 10 | 26.64 | 26.64 | R | MPC | 112-39-0
|  | Styrene-acrylonitrile Trimer | 26.93 | 10.89 | 25 | 10 | 4.356 | 4.356 | R | TIC | 57964-39-3
|  | Unknown | 27.13 | 22.31 | 25 | 10 | 8.924 | 8.924 | R | U | -
|  | Oleanitrile | 28.02 | 9.68 | 25 | 10 | 3.872 | 3.872 | R | MPC | 112-91-4
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.1 | 69.49 | 25 | 10 | 27.796 | 27.796 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.15 | 124.07 | 25 | 10 | 49.628 | 49.628 | R | U | -
|  | Octadecanoic Acid methyl ester | 28.3 | 112.07 | 25 | 10 | 44.828 | 44.828 | R | MPC | 112-61-8
|  | Oleic Acid | 28.44 | 21.82 | 25 | 10 | 8.728 | 8.728 | R | TIC | 112-80-1
|  | Octadecanoic Acid | 28.58 | 24.82 | 25 | 10 | 9.928 | 9.928 | R | MPC | 57-11-4
|  | Unknown | 29.56 | 34.23 | 25 | 10 | 13.692 | 13.692 | R | U | -
|  | Unknown | 29.69 | 18.27 | 25 | 10 | 7.308 | 7.308 | R | U | -
| Accu-Chek d10 S3 | 1,4-Butanediol | 10.43 | 12.93 | 25 | 10 | 5.172 | 5.172 | R | MPC | 110-63-4
|  | 1-[4-(1-methylethyl)phenyl]-Ethanone | 18.08 | 10.13 | 25 | 10 | 4.052 | 4.052 | R | MPC | 645-13-6
|  | 1,6-Dioxacyclododecane-7,12-dione | 22.05 | 85.32 | 25 | 10 | 34.128 | 34.128 | R | MPC | 777-95-7
|  | Dodecyl acrylate | 23.85 | 30.1 | 25 | 10 | 12.04 | 12.04 | R | TIC | 2156-97-0
|  | (1-hydroxycyclohexyl)phenyl-Methanone | 24.19 | 19.73 | 25 | 10 | 7.892 | 7.892 | R | MPC | 947-19-3
|  | Unknown | 24.3 | 14.35 | 25 | 10 | 5.74 | 5.74 | R | U | -
|  | Hexadecanoic acid, methyl ester | 26.48 | 69.47 | 25 | 10 | 27.788 | 27.788 | R | MPC | 112-39-0
|  | Styrene-acrylonitrile Trimer | 26.93 | 8.88 | 25 | 10 | 3.552 | 3.552 | R | TIC | 57964-39-3
|  | Unknown | 27.13 | 18.21 | 25 | 10 | 7.284 | 7.284 | R | U | -
|  | Oleanitrile | 28.02 | 10.13 | 25 | 10 | 4.052 | 4.052 | R | MPC | 112-91-4
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.1 | 70.81 | 25 | 10 | 28.324 | 28.324 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.15 | 130.19 | 25 | 10 | 52.076 | 52.076 | R | U | -
|  | Octadecanoic Acid methyl ester | 28.3 | 114.84 | 25 | 10 | 45.936 | 45.936 | R | MPC | 112-61-8
|  | Oleic Acid | 28.44 | 21.97 | 25 | 10 | 8.788 | 8.788 | R | TIC | 112-80-1
|  | Octadecanoic Acid | 28.58 | 20.42 | 25 | 10 | 8.168 | 8.168 | R | MPC | 57-11-4
|  | Unknown | 29.58 | 30.15 | 25 | 10 | 12.06 | 12.06 | R | U | -
|  | Unknown | 29.69 | 16.81 | 25 | 10 | 6.724 | 6.724 | R | U | -
| Neria d10 S1 | 1,4-Butanediol | 10.45 | 12.89 | 25 | 10 | 5.156 | 5.156 | R | MPC | 110-63-4
|  | 1-[4-(1-methylethyl)phenyl]-Ethanone | 18.08 | 10.96 | 25 | 10 | 4.384 | 4.384 | R | MPC | 645-13-6
|  | Butylated Hydroxytoluene | 21.18 | 18.33 | 25 | 10 | 7.332 | 7.332 | R | MPC | 128-37-0
|  | 1,6-Dioxacyclododecane-7,12-dione | 22.05 | 112.99 | 25 | 10 | 45.196 | 45.196 | R | MPC | 777-95-7
|  | Dodecanoic acid, 1-methylethyl ester | 22.88 | 15.72 | 25 | 10 | 6.288 | 6.288 | R | MPC | 10233-13-3
|  | Dodecyl acrylate | 23.85 | 32.95 | 25 | 10 | 13.18 | 13.18 | R | TIC | 2156-97-0
|  | (1-hydroxycyclohexyl)phenyl-Methanone | 24.19 | 33.43 | 25 | 10 | 13.372 | 13.372 | R | MPC | 947-19-3
|  | Unknown | 24.3 | 16.48 | 25 | 10 | 6.592 | 6.592 | R | U | -
|  | Hexadecanoic acid, methyl ester | 26.48 | 62.97 | 25 | 10 | 25.188 | 25.188 | R | MPC | 112-39-0
|  | n-Hexadecanoic Acid | 26.8 | 18.02 | 25 | 10 | 7.208 | 7.208 | R | MPC | 57-10-3
|  | Unknown | 27.13 | 20.53 | 25 | 10 | 8.212 | 8.212 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.1 | 66.62 | 25 | 10 | 26.648 | 26.648 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.15 | 115.72 | 25 | 10 | 46.288 | 46.288 | R | U | -
|  | Octadecanoic Acid methyl ester | 28.3 | 104.85 | 25 | 10 | 41.94 | 41.94 | R | MPC | 112-61-8
|  | Oleic Acid | 28.44 | 45.47 | 25 | 10 | 18.188 | 18.188 | R | TIC | 112-80-1
|  | Octadecanoic Acid | 28.58 | 41.9 | 25 | 10 | 16.76 | 16.76 | R | MPC | 57-11-4
|  | Unknown | 29.58 | 27.81 | 25 | 10 | 11.124 | 11.124 | R | U | -
|  | Unknown | 29.69 | 14.19 | 25 | 10 | 5.676 | 5.676 | R | U | -
| Neria d10 S2 | 1,4-Butanediol | 10.45 | 12.19 | 25 | 10 | 4.876 | 4.876 | R | MPC | 110-63-4
|  | 1-[4-(1-methylethyl)phenyl]-Ethanone | 18.08 | 10.56 | 25 | 10 | 4.224 | 4.224 | R | MPC | 645-13-6
|  | Butylated Hydroxytoluene | 21.18 | 19.16 | 25 | 10 | 7.664 | 7.664 | R | MPC | 128-37-0
|  | 1,6-Dioxacyclododecane-7,12-dione | 22.05 | 112.51 | 25 | 10 | 45.004 | 45.004 | R | MPC | 777-95-7
|  | Dodecanoic acid, 1-methylethyl ester | 22.88 | 11.76 | 25 | 10 | 4.704 | 4.704 | R | MPC | 10233-13-3
|  | Dodecyl acrylate | 23.85 | 33.07 | 25 | 10 | 13.228 | 13.228 | R | TIC | 2156-97-0
|  | (1-hydroxycyclohexyl)phenyl-Methanone | 24.19 | 31.9 | 25 | 10 | 12.76 | 12.76 | R | MPC | 947-19-3
|  | Unknown | 24.3 | 16.79 | 25 | 10 | 6.716 | 6.716 | R | U | -
|  | Hexadecanoic acid, methyl ester | 26.48 | 62.82 | 25 | 10 | 25.128 | 25.128 | R | MPC | 112-39-0
|  | n-Hexadecanoic Acid | 26.79 | 17.13 | 25 | 10 | 6.852 | 6.852 | R | MPC | 57-10-3
|  | Unknown | 27.13 | 21.01 | 25 | 10 | 8.404 | 8.404 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.1 | 67.61 | 25 | 10 | 27.044 | 27.044 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.15 | 113.48 | 25 | 10 | 45.392 | 45.392 | R | U | -
|  | Octadecanoic Acid methyl ester | 28.3 | 106.09 | 25 | 10 | 42.436 | 42.436 | R | MPC | 112-61-8
|  | Oleic Acid | 28.44 | 43.13 | 25 | 10 | 17.252 | 17.252 | R | TIC | 112-80-1
|  | Octadecanoic Acid | 28.59 | 40.29 | 25 | 10 | 16.116 | 16.116 | R | MPC | 57-11-4
|  | Unknown | 29.58 | 28.05 | 25 | 10 | 11.22 | 11.22 | R | U | -
|  | Unknown | 29.69 | 14.07 | 25 | 10 | 5.628 | 5.628 | R | U | -
| Neria d10 S3 | 1,4-Butanediol | 10.45 | 13.11 | 25 | 10 | 5.244 | 5.244 | R | MPC | 110-63-4
|  | 1-[4-(1-methylethyl)phenyl]-Ethanone | 18.08 | 10.25 | 25 | 10 | 4.1 | 4.1 | R | MPC | 645-13-6
|  | Butylated Hydroxytoluene | 21.18 | 19.41 | 25 | 10 | 7.764 | 7.764 | R | MPC | 128-37-0
|  | 1,6-Dioxacyclododecane-7,12-dione | 22.05 | 114.94 | 25 | 10 | 45.976 | 45.976 | R | MPC | 777-95-7
|  | Dodecanoic acid, 1-methylethyl ester | 22.88 | 17.01 | 25 | 10 | 6.804 | 6.804 | R | MPC | 10233-13-3
|  | Dodecyl acrylate | 23.85 | 32.25 | 25 | 10 | 12.9 | 12.9 | R | TIC | 2156-97-0
|  | (1-hydroxycyclohexyl)phenyl-Methanone | 24.19 | 34.53 | 25 | 10 | 13.812 | 13.812 | R | MPC | 947-19-3
|  | Unknown | 24.3 | 16.21 | 25 | 10 | 6.484 | 6.484 | R | U | -
|  | Hexadecanoic acid, methyl ester | 26.48 | 63.16 | 25 | 10 | 25.264 | 25.264 | R | MPC | 112-39-0
|  | n-Hexadecanoic Acid | 26.79 | 17.22 | 25 | 10 | 6.888 | 6.888 | R | MPC | 57-10-3
|  | Unknown | 27.13 | 20.45 | 25 | 10 | 8.18 | 8.18 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.1 | 68.29 | 25 | 10 | 27.316 | 27.316 | R | U | -
|  | (7 or 9)-Octadecanoic Acid methyl ester | 28.15 | 115.32 | 25 | 10 | 46.128 | 46.128 | R | U | -
|  | Octadecanoic Acid methyl ester | 28.3 | 106.69 | 25 | 10 | 42.676 | 42.676 | R | MPC | 112-61-8
|  | Oleic Acid | 28.44 | 43.33 | 25 | 10 | 17.332 | 17.332 | R | TIC | 112-80-1
|  | Octadecanoic Acid | 28.59 | 40.75 | 25 | 10 | 16.3 | 16.3 | R | MPC | 57-11-4
|  | Unknown | 29.56 | 26.86 | 25 | 10 | 10.744 | 10.744 | R | U | -
|  | Unknown | 29.69 | 13.53 | 25 | 10 | 5.412 | 5.412 | R | U | -
[8 empty] | Blank | Reportable
Drug solution 1week50C |  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
1512020oe |  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
|  | unknown | 31.54 | 6 | 5 | 10 | 12 | 12 | NR | N/A | N/A
| mb1511020OE | unknown | 32.08 | 68.61 | 5 | 10 | 137.22 | 137.22 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 21.7 | 5 | 10 | 43.4 | 43.4 | NR | MPC | 58-00-4
| control | unknown | 32.08 | 47.61 | 5 | 10 | 95.22 | -42 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 18.91 | 5 | 10 | 37.82 | -5.58 | NR | MPC | 58-00-4
| S1(Accu-Chek) | 1,6-Dioxacyclododecane-7,12-dione | 22.02 | 9.79 | 5 | 10 | 19.58 | 19.58 | R | MPC | 777-95-7
|  | unknown | 31.54 | 11.69 | 5 | 10 | 23.38 | 11.38 | NR | N/A | N/A
|  | unknown | 32.08 | 42.71 | 5 | 10 | 85.42 | -51.8 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 13.91 | 5 | 10 | 27.82 | -15.58 | NR | MPC | 58-00-4
| S2(Accu-Chek) | 1,6-Dioxacyclododecane-7,12-dione | 22.02 | 10.54 | 5 | 10 | 21.08 | 21.08 | R | MPC | 777-95-7
|  | unknown | 31.54 | 12.95 | 5 | 10 | 25.9 | 13.9 | NR | N/A | N/A
|  | unknown | 32.08 | 43.64 | 5 | 10 | 87.28 | -49.94 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 16.08 | 5 | 10 | 32.16 | -11.24 | NR | MPC | 58-00-4
| S3(Accu-Chek) | 1,6-Dioxacyclododecane-7,12-dione | 22.02 | 8.79 | 5 | 10 | 17.58 | 17.58 | R | MPC | 777-95-7
|  | unknown | 31.54 | 15.98 | 5 | 10 | 31.96 | 19.96 | NR | N/A | N/A
|  | unknown | 32.08 | 36.71 | 5 | 10 | 73.42 | -63.8 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 11.27 | 5 | 10 | 22.54 | -20.86 | NR | MPC | 58-00-4
| S1(Neria) | 1,6-Dioxacyclododecane-7,12-dione | 22.02 | 12.94 | 5 | 10 | 25.88 | 25.88 | R | MPC | 777-95-7
|  | unknown | 31.54 | 17.54 | 5 | 10 | 35.08 | 23.08 | NR | N/A | N/A
|  | unknown | 32.08 | 34.89 | 5 | 10 | 69.78 | -67.44 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 7.45 | 5 | 10 | 14.9 | -28.5 | NR | MPC | 58-00-4
| S2(Neria) | 1,6-Dioxacyclododecane-7,12-dione | 22.02 | 14.47 | 5 | 10 | 28.94 | 28.94 | R | MPC | 777-95-7
|  | unknown | 31.54 | 18.47 | 5 | 10 | 36.94 | 24.94 | NR | N/A | N/A
|  | unknown | 32.08 | 40.44 | 5 | 10 | 80.88 | -56.34 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.99 | 8.89 | 5 | 10 | 17.78 | -25.62 | NR | MPC | 58-00-4
| S3(Neria) | 1,6-Dioxacyclododecane-7,12-dione | 22.02 | 14.56 | 5 | 10 | 29.12 | 29.12 | R | MPC | 777-95-7
|  | unknown | 31.53 | 1.34 | 5 | 10 | 2.68 | -9.32 | NR | N/A | N/A
|  | unknown | 32.08 | 53.26 | 5 | 10 | 106.52 | -30.7 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.98 | 34.31 | 5 | 10 | 68.62 | 25.22 | NR | MPC | 58-00-4
| S1(Syringe) | unknown | 32.08 | 41.66 | 5 | 10 | 83.32 | -53.9 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.98 | 33.84 | 5 | 10 | 67.68 | 24.28 | NR | MPC | 58-00-4
| S2(Syringe) | unknown | 32.08 | 35.09 | 5 | 10 | 70.18 | -67.04 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.98 | 39.88 | 5 | 10 | 79.76 | 36.36 | NR | MPC | 58-00-4
| S2(Syringe) | unknown | 32.08 | 39.73 | 5 | 10 | 79.46 | -57.76 | NR | N/A | N/A
|  | 4H-Dibenzo[de,g]quinoline-10,11-diol,5,6,6a,7-tetrahydro-6-methyl-, (6aR)- | 32.98 | 30.67 | 5 | 10 | 61.34 | 17.94 | NR | MPC | 58-00-4
SVOC
[8 empty] | Blank | Reportable
IPA 1week 50C |  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
| S1(Syringe) | Ethane, 1,1,2,2-tetrachloro- | 9.9 | 17.16 | 25 | 10 | 6.864 | 6.864 | NR | MPC | 79-34-5
|  | unknown | 18.63 | 15.25 | 25 | 10 | 6.1 | 6.1 | NR | N/A | N/A
|  | unknown | 20.37 | 8.51 | 25 | 10 | 3.404 | 3.404 | NR | N/A | N/A
|  | Butylated Hydroxytoluene | 21.14 | 101.76 | 25 | 10 | 40.704 | 40.704 | NR | MPC | 128-37-0
|  | 1,7-Octadiene, 2-methyl-6-methylene- | 22.12 | 22.17 | 25 | 10 | 8.868 | 8.868 | NR | TIC | 1686-30-2
|  | n-Hexadecanoic acid | 26.77 | 20.46 | 25 | 10 | 8.184 | 8.184 | NR | MPC | 57-10-3
|  | Oleic acid | 28.39 | 64.4 | 25 | 10 | 25.76 | 25.76 | NR | MPC | 112-80-1
|  | Octadecanoic acid | 28.57 | 61.1 | 25 | 10 | 24.44 | 24.44 | NR | MPC | 57-11-4
|  | 2-hydroxy-1-(hydroxymethyl)ethyl ester-9-(Z)-octadecenoic acid | 30.56 | 22.19 | 25 | 10 | 8.876 | 8.876 | NR | MPC | 3443-84-3
|  | unknown phenol | 30.71 | 22.65 | 25 | 10 | 9.06 | 9.06 | NR | N/A | N/A
|  | unknown phenol | 30.81 | 13.57 | 25 | 10 | 5.428 | 5.428 | NR | N/A | N/A
|  | unknown phenol | 31.2 | 14.08 | 25 | 10 | 5.632 | 5.632 | NR | N/A | N/A
|  | unknown | 36.5 | 13.02 | 25 | 10 | 5.208 | 5.208 | NR | N/A | N/A
|  | unknown | 36.74 | 15.53 | 25 | 10 | 6.212 | 6.212 | NR | N/A | N/A
|  | unknown | 45.19 | 22.08 | 25 | 10 | 8.832 | 8.832 | NR | N/A | N/A
|  | unknown | 50.56 | 24.29 | 25 | 10 | 9.716 | 9.716 | NR | N/A | N/A
|  | unknown | 52.56 | 19.28 | 25 | 10 | 7.712 | 7.712 | NR | N/A | N/A
|  | unknown | 57.01 | 81.52 | 25 | 10 | 32.608 | 32.608 | NR | N/A | N/A
| S2(Syringe) | Ethane, 1,1,2,2-tetrachloro- | 9.9 | 15.12 | 25 | 10 | 6.048 | 6.048 | NR | MPC | 79-34-5
|  | unknown | 18.62 | 13.36 | 25 | 10 | 5.344 | 5.344 | NR | N/A | N/A
|  | unknown | 20.37 | 10.03 | 25 | 10 | 4.012 | 4.012 | NR | N/A | N/A
|  | Butylated Hydroxytoluene | 21.14 | 93.68 | 25 | 10 | 37.472 | 37.472 | NR | MPC | 128-37-0
|  | 1,7-Octadiene, 2-methyl-6-methylene- | 22.12 | 20.36 | 25 | 10 | 8.144 | 8.144 | NR | TIC | 1686-30-2
|  | n-Hexadecanoic acid | 26.77 | 22.27 | 25 | 10 | 8.908 | 8.908 | NR | MPC | 57-10-3
|  | Oleic acid | 28.39 | 65.7 | 25 | 10 | 26.28 | 26.28 | NR | MPC | 112-80-1
|  | Octadecanoic acid | 28.57 | 58.61 | 25 | 10 | 23.444 | 23.444 | NR | MPC | 57-11-4
|  | 2-hydroxy-1-(hydroxymethyl)ethyl ester-9-(Z)-octadecenoic acid | 30.56 | 21.05 | 25 | 10 | 8.42 | 8.42 | NR | MPC | 3443-84-3
|  | unknown phenol | 30.71 | 21.53 | 25 | 10 | 8.612 | 8.612 | NR | N/A | N/A
|  | unknown phenol | 30.81 | 13.53 | 25 | 10 | 5.412 | 5.412 | NR | N/A | N/A
|  | unknown phenol | 31.19 | 13.41 | 25 | 10 | 5.364 | 5.364 | NR | N/A | N/A
|  | unknown | 36.5 | 13.74 | 25 | 10 | 5.496 | 5.496 | NR | N/A | N/A
|  | unknown | 36.74 | 18.01 | 25 | 10 | 7.204 | 7.204 | NR | N/A | N/A
|  | unknown | 45.17 | 21.4 | 25 | 10 | 8.56 | 8.56 | NR | N/A | N/A
|  | unknown | 50.53 | 20.37 | 25 | 10 | 8.148 | 8.148 | NR | N/A | N/A
|  | unknown | 52.55 | 17.72 | 25 | 10 | 7.088 | 7.088 | NR | N/A | N/A
|  | unknown | 56.99 | 80.4 | 25 | 10 | 32.16 | 32.16 | NR | N/A | N/A
| S3(Syringe) | Ethane, 1,1,2,2-tetrachloro- | 9.9 | 15.9 | 25 | 10 | 6.36 | 6.36 | NR | MPC | 79-34-5
|  | unknown | 18.62 | 13.7 | 25 | 10 | 5.48 | 5.48 | NR | N/A | N/A
|  | unknown | 20.37 | 8.81 | 25 | 10 | 3.524 | 3.524 | NR | N/A | N/A
|  | Butylated Hydroxytoluene | 21.14 | 99.22 | 25 | 10 | 39.688 | 39.688 | NR | MPC | 128-37-0
|  | 1,7-Octadiene, 2-methyl-6-methylene- | 22.12 | 20.82 | 25 | 10 | 8.328 | 8.328 | NR | TIC | 1686-30-2
|  | n-Hexadecanoic acid | 26.77 | 23.39 | 25 | 10 | 9.356 | 9.356 | NR | MPC | 57-10-3
|  | Oleic acid | 28.39 | 62.9 | 25 | 10 | 25.16 | 25.16 | NR | MPC | 112-80-1
|  | Octadecanoic acid | 28.57 | 55.81 | 25 | 10 | 22.324 | 22.324 | NR | MPC | 57-11-4
|  | 2-hydroxy-1-(hydroxymethyl)ethyl ester-9-(Z)-octadecenoic acid | 30.56 | 19.58 | 25 | 10 | 7.832 | 7.832 | NR | MPC | 3443-84-3
|  | unknown phenol | 30.71 | 20.65 | 25 | 10 | 8.26 | 8.26 | NR | N/A | N/A
|  | unknown phenol | 30.81 | 12.85 | 25 | 10 | 5.14 | 5.14 | NR | N/A | N/A
|  | unknown phenol | 31.19 | 13.51 | 25 | 10 | 5.404 | 5.404 | NR | N/A | N/A
|  | unknown | 36.5 | 13.35 | 25 | 10 | 5.34 | 5.34 | NR | N/A | N/A
|  | unknown | 36.74 | 15.07 | 25 | 10 | 6.028 | 6.028 | NR | N/A | N/A
|  | unknown | 45.17 | 20.95 | 25 | 10 | 8.38 | 8.38 | NR | N/A | N/A
|  | unknown | 50.56 | 21.03 | 25 | 10 | 8.412 | 8.412 | NR | N/A | N/A
|  | unknown | 52.52 | 19.09 | 25 | 10 | 7.636 | 7.636 | NR | N/A | N/A
|  | unknown | 56.99 | 77.73 | 25 | 10 | 31.092 | 31.092 | NR | N/A | N/A
Compounds observed in blank not reported, as they are not test article/control related.
Corrected Result (ug/mL) = (Instrument Result (ug/mL) * Dilution Factor)/Concentration Factor
Blank/Control Corrected Result (ug/mL) = Correct Result (ug/mL) - Corrected Result Corrected for Blank/Control (ug/mL)
Concentration Factor = Initial Extraction Volume/Concentrated Volume
MPC - >80%
TIC - <80%

## NVOC
15-03379-N1 |  | NVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
Acidified Water | mb15110250oe | Palmitic Acid | 5.892 | 5.37 | 25 | 1 | 0.2148 | 0.2148 | - | IC | 57-10-3
|  | Oleic Acid | 6.259 | 0.23 | 25 | 1 | 0.0092 | 0.0092 | - | IC | 112-80-1
|  | Stearic Acid | 7.08 | 4.81 | 25 | 1 | 0.1924 | 0.1924 | - | IC | 57-11-4
|  | Erucamide | 7.184 | 1.72 | 25 | 1 | 0.0688 | 0.0688 | - | IC | 112-84-5
|  | Palmitic Acid | 5.881 | 3.67 | 25 | 1 | 0.1468 | -0.068 | NR | IC | 57-10-3
| Control | Stearic Acid | 7.055 | 3.75 | 25 | 1 | 0.15 | -0.0424 | NR | IC | 57-11-4
|  | Erucamide | 7.172 | 2.11 | 25 | 1 | 0.0844 | 0.0156 | NR | IC | 112-84-5
|  | Oleamide | 4.936 | 1.49 | 25 | 1 | 0.0596 | 0.0596 | R | IC | 301-02-0
|  | Palmitic Acid | 5.822 | 3.71 | 25 | 1 | 0.1484 | -0.0664 | NR | IC | 57-10-3
| Accu-Chek S1 | Oleic Acid | 6.187 | 0.26 | 25 | 1 | 0.0104 | 0.0012 | NR | IC | 112-80-1
|  | Stearic Acid | 7.024 | 3.63 | 25 | 1 | 0.1452 | -0.0472 | NR | IC | 57-11-4
|  | Erucamide | 7.136 | 2 | 25 | 1 | 0.08 | 0.0112 | NR | IC | 112-84-5
|  | Oleamide | 4.932 | 1.68 | 25 | 1 | 0.0672 | 0.0672 | R | IC | 301-02-0
|  | Palmitic Acid | 5.844 | 3.6 | 25 | 1 | 0.144 | -0.0708 | NR | IC | 57-10-3
| Accu-Chek S2 | Oleic Acid | 6.183 | 0.35 | 25 | 1 | 0.014 | 0.0048 | NR | IC | 112-80-1
|  | Stearic Acid | 7.047 | 3.133 | 25 | 1 | 0.12532 | -0.06708 | NR | IC | 57-11-4
|  | Erucamide | 7.149 | 2.91 | 25 | 1 | 0.1164 | 0.0476 | NR | IC | 112-84-5
|  | Oleamide | 4.927 | 1.86 | 25 | 1 | 0.0744 | 0.0744 | R | IC | 301-02-0
|  | Palmitic Acid | 5.825 | 4.56 | 25 | 1 | 0.1824 | -0.0324 | NR | IC | 57-10-3
| Accu-Chek S3 | Oleic Acid | 6.232 | 0.43 | 25 | 1 | 0.0172 | 0.008 | NR | IC | 112-80-1
|  | Stearic Acid | 7.041 | 3.82 | 25 | 1 | 0.1528 | -0.0396 | NR | IC | 57-11-4
|  | Erucamide | 7.138 | 3.22 | 25 | 1 | 0.1288 | 0.06 | NR | IC | 112-84-5
|  | Oleamide | 4.958 | 2.23 | 25 | 1 | 0.0892 | 0.0892 | R | IC | 301-02-0
|  | Palmitic Acid | 5.831 | 4.33 | 25 | 1 | 0.1732 | -0.0416 | NR | IC | 57-10-3
| Neria S1 | Oleic Acid | 6.213 | 0.54 | 25 | 1 | 0.0216 | 0.0124 | NR | IC | 112-80-1
|  | Stearic Acid | 7.052 | 3.22 | 25 | 1 | 0.1288 | -0.0636 | NR | IC | 57-11-4
|  | Erucamide | 7.142 | 4.31 | 25 | 1 | 0.1724 | 0.1036 | NR | IC | 112-84-5
|  | Oleamide | 4.734 | 1.76 | 25 | 1 | 0.0704 | 0.0704 | R | IC | 301-02-0
|  | Palmitic Acid | 5.611 | 4.28 | 25 | 1 | 0.1712 | -0.0436 | NR | IC | 57-10-3
| Neria S2 | Stearic Acid | 6.772 | 4.18 | 25 | 1 | 0.1672 | -0.0252 | NR | IC | 57-11-4
|  | Erucamide | 6.877 | 2.87 | 25 | 1 | 0.1148 | 0.046 | NR | IC | 112-84-5
|  | Oleamide | 4.745 | 1.65 | 25 | 1 | 0.066 | 0.066 | R | IC | 301-02-0
|  | Palmitic Acid | 5.595 | 3.84 | 25 | 1 | 0.1536 | -0.0612 | NR | IC | 57-10-3
| Neria S3 | Oleic Acid | 6 | 0.33 | 25 | 1 | 0.0132 | 0.004 | NR | IC | 112-80-1
|  | Stearic Acid | 6.767 | 3.64 | 25 | 1 | 0.1456 | -0.0468 | NR | IC | 57-11-4
|  | Erucamide | 6.894 | 2.7 | 25 | 1 | 0.108 | 0.0392 | NR | IC | 112-84-5
|  | Oleamide | 4.772 | 2.31 | 25 | 1 | 0.0924 | 0.0924 | R | IC | 301-02-0
|  | Palmitic Acid | 5.614 | 4.17 | 25 | 1 | 0.1668 | -0.048 | NR | IC | 57-10-3
| Syringe S1 | Oleic Acid | 5.971 | 0.56 | 25 | 1 | 0.0224 | 0.0132 | NR | IC | 112-80-1
|  | Stearic Acid | 6.804 | 2.82 | 25 | 1 | 0.1128 | -0.0796 | NR | IC | 57-11-4
|  | Erucamide | 6.888 | 4.47 | 25 | 1 | 0.1788 | 0.11 | NR | IC | 112-84-5
|  | Irganox 1010 | 9.635 | 0 | 25 | 1 | 0 | 0 | NR | IC | 98584-37-3
|  | Oleamide | 4.75 | 1.22 | 25 | 1 | 0.0488 | 0.0488 | R | IC | 301-02-0
|  | Palmitic Acid | 5.603 | 3.37 | 25 | 1 | 0.1348 | -0.08 | NR | IC | 57-10-3
| Syringe S2 | Oleic Acid | 5.993 | 0.32 | 25 | 1 | 0.0128 | 0.0036 | NR | IC | 112-80-1
|  | Stearic Acid | 6.786 | 3.47 | 25 | 1 | 0.1388 | -0.0536 | NR | IC | 57-11-4
|  | Erucamide | 6.886 | 2.19 | 25 | 1 | 0.0876 | 0.0188 | NR | IC | 112-84-5
|  | Irganox 1010 | 9.64 | 0 | 25 | 1 | 0 | 0 | NR | IC | 98584-37-3
|  | Oleamide | 4.768 | 1.68 | 25 | 1 | 0.0672 | 0.0672 | R | IC | 301-02-0
|  | Palmitic Acid | 5.603 | 5.33 | 25 | 1 | 0.2132 | -0.0016 | NR | IC | 57-10-3
| Syringe S3 | Oleic Acid | 5.965 | 0.59 | 25 | 1 | 0.0236 | 0.0144 | NR | IC | 112-80-1
|  | Stearic Acid | 6.807 | 4.03 | 25 | 1 | 0.1612 | -0.0312 | NR | IC | 57-11-4
|  | Erucamide | 6.886 | 4.94 | 25 | 1 | 0.1976 | 0.1288 | NR | IC | 112-84-5
|  | Irganox 1010 | 9.63 | 0 | 25 | 1 | 0 | 0 | NR | IC | 98584-37-3
BRL - Below reporting Limit (<1 ug/mL)
Compounds observed in blank not reported, as they are not test article/control related.
Corrected Resulted (ug/mL) = (Instrument Result (ug/mL) * Dilution Factor)/Concentration Factor
Blank/Control Corrected Result (ug/mL) = Correct Result (ug/mL) - Corrected Result Corrected for Blank/Control (ug/mL)
Concentration Factor = Initial Extraction Volume/Concentrated Volume
Highlighted cells are peaks that were integrated but not automatically reported. Need to be analyst verified.
15-03379-N1 |  | NVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
Drug Product | mb1511018oe | Oleamide | 5.028 | 1.16 | 1 | 5 | 5.8 | 5.8 | - | IC | 301-02-0
|  | Palmitic Acid | 5.892 | 0.77 | 1 | 5 | 3.85 | 3.85 | - | IC | 57-10-3
|  | Oleic Acid | 6.256 | 0.29 | 1 | 5 | 1.45 | 1.45 | - | IC | 112-80-1
|  | Stearic Acid | 7.1 | 0.48 | 1 | 5 | 2.4 | 2.4 | - | IC | 57-11-4
|  | Erucamide | 7.208 | 1.14 | 1 | 5 | 5.7 | 5.7 | - | IC | 112-84-5
| Control | Palmitic Acid | 5.835 | 1.13 | 1 | 5 | 5.65 | 1.8 | NR | IC | 57-10-3
|  | Oleic Acid | 6.192 | 1.03 | 1 | 5 | 5.15 | 3.7 | NR | IC | 112-80-1
|  | Stearic Acid | 7.044 | 1.08 | 1 | 5 | 5.4 | 3 | NR | IC | 57-11-4
|  | Erucamide | 7.146 | 1.09 | 1 | 5 | 5.45 | -0.25 | NR | IC | 112-84-5
| Accu-Chek S1 | Palmitic Acid | 5.835 | 2.01 | 1 | 5 | 10.05 | 6.2 | NR | IC | 57-10-3
|  | Oleic Acid | 6.218 | 0.96 | 1 | 5 | 4.8 | 3.35 | NR | IC | 112-80-1
|  | Stearic Acid | 7.053 | 2.07 | 1 | 5 | 10.35 | 7.95 | NR | IC | 57-11-4
|  | Erucamide | 7.178 | 1.33 | 1 | 5 | 6.65 | 0.95 | NR | IC | 112-84-5
| Accu-Chek S2 | Palmitic Acid | 5.855 | 1.27 | 1 | 5 | 6.35 | 2.5 | NR | IC | 57-10-3
|  | Oleic Acid | 6.211 | 0.71 | 1 | 5 | 3.55 | 2.1 | NR | IC | 112-80-1
|  | Stearic Acid | 7.049 | 1.62 | 1 | 5 | 8.1 | 5.7 | NR | IC | 57-11-4
|  | Erucamide | 7.149 | 1.23 | 1 | 5 | 6.15 | 0.45 | NR | IC | 112-84-5
| Accu-Chek S3 | Palmitic Acid | 5.858 | 0.66 | 1 | 5 | 3.3 | -0.55 | NR | IC | 57-10-3
|  | Oleic Acid | 6.218 | 0.35 | 1 | 5 | 1.75 | 0.3 | NR | IC | 112-80-1
|  | Stearic Acid | 7.053 | 1.19 | 1 | 5 | 5.95 | 3.55 | NR | IC | 57-11-4
|  | Erucamide | 7.169 | 1.1 | 1 | 5 | 5.5 | -0.2 | NR | IC | 112-84-5
| Neria S1 | Palmitic Acid | 5.856 | 1.12 | 1 | 5 | 5.6 | 1.75 | NR | IC | 57-10-3
|  | Oleic Acid | 6.244 | 0.29 | 1 | 5 | 1.45 | 0 | NR | IC | 112-80-1
|  | Stearic Acid | 7.058 | 1.28 | 1 | 5 | 6.4 | 4 | NR | IC | 57-11-4
|  | Erucamide | 7.172 | 1.19 | 1 | 5 | 5.95 | 0.25 | NR | IC | 112-84-5
| Neria S2 | Oleamide | 4.996 | 1.16 | 1 | 5 | 5.8 | 0 | NR | IC | 301-02-0
|  | Palmitic Acid | 5.879 | 1.36 | 1 | 5 | 6.8 | 2.95 | NR | IC | 57-10-3
|  | Oleic Acid | 6.226 | 0.65 | 1 | 5 | 3.25 | 1.8 | NR | IC | 112-80-1
|  | Stearic Acid | 7.064 | 1.5 | 1 | 5 | 7.5 | 5.1 | NR | IC | 57-11-4
|  | Erucamide | 7.182 | 1.52 | 1 | 5 | 7.6 | 1.9 | NR | IC | 112-84-5
| Neria S3 | Palmitic Acid | 5.827 | 0.71 | 1 | 5 | 3.55 | -0.3 | NR | IC | 57-10-3
|  | Oleic Acid | 6.201 | 3.98 | 1 | 5 | 19.9 | 18.45 | NR | IC | 112-80-1
|  | Stearic Acid | 7.047 | 0.7 | 1 | 5 | 3.5 | 1.1 | NR | IC | 57-11-4
|  | Erucamide | 7.186 | 1.13 | 1 | 5 | 5.65 | -0.05 | NR | IC | 112-84-5
| Syringe S1 | Palmitic Acid | 5.882 | 2.02 | 1 | 5 | 10.1 | 6.25 | NR | IC | 57-10-3
|  | Oleic Acid | 6.24 | 1.51 | 1 | 5 | 7.55 | 6.1 | NR | IC | 112-80-1
|  | Stearic Acid | 7.081 | 3.23 | 1 | 5 | 16.15 | 13.75 | NR | IC | 57-11-4
|  | Erucamide | 7.187 | 1.16 | 1 | 5 | 5.8 | 0.1 | NR | IC | 112-84-5
| Syringe S2 | Palmitic Acid | 5.875 | 1.65 | 1 | 5 | 8.25 | 4.4 | NR | IC | 57-10-3
|  | Oleic Acid | 6.251 | 0.94 | 1 | 5 | 4.7 | 3.25 | NR | IC | 112-80-1
|  | Stearic Acid | 7.091 | 2.22 | 1 | 5 | 11.1 | 8.7 | NR | IC | 57-11-4
|  | Erucamide | 7.199 | 1.26 | 1 | 5 | 6.3 | 0.6 | NR | IC | 112-84-5
| Syringe S3 | Palmitic Acid | 5.881 | 3.35 | 1 | 5 | 16.75 | 12.9 | NR | IC | 57-10-3
|  | Oleic Acid | 6.249 | 3.28 | 1 | 5 | 16.4 | 14.95 | NR | IC | 112-80-1
|  | Stearic Acid | 7.089 | 4.37 | 1 | 5 | 21.85 | 19.45 | NR | IC | 57-11-4
|  | Erucamide | 7.201 | 1.12 | 1 | 5 | 5.6 | -0.1 | NR | IC | 112-84-5
BRL - Below reporting Limit (<1 ug/mL)
Compounds observed in blank not reported, as they are not test article/control related.
Corrected Resulted (ug/mL) = (Instrument Result (ug/mL) * Dilution Factor)/Concentration Factor
Blank/Control Corrected Result (ug/mL) = Correct Result (ug/mL) - Corrected Result Corrected for Blank/Control (ug/mL)
Concentration Factor = Initial Extraction Volume/Concentrated Volume
Highlighted cells are peaks that were integrated but not automatically reported. Need to be analyst verified.
15-03379-N1 |  | NVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
IPA | mb1511024oe | None | N/A | 0 | 25 | 10 | 0 | 0 | - | - | -
| Control | Palmitic Acid | 5.949 | 0.59 | 25 | 10 | 0.236 | 0.236 | NR | IC | 57-10-3
|  | Oleic Acid | 6.374 | 0.29 | 25 | 10 | 0.116 | 0.116 | NR | IC | 112-80-1
|  | Stearic Acid | 7.163 | 1.24 | 25 | 10 | 0.496 | 0.496 | NR | IC | 57-11-4
|  | Irganox 1076 | 11.386 | 0.28 | 25 | 10 | 0.112 | 0.112 | NR | IC | 2082-79-3
| Accu-Chek S1 | BHT | 4.102 | 8.12 | 25 | 10 | 3.248 | 3.248 | R | IC | 128-37-0
|  | Myristic Acid | 4.517 | 1.28 | 25 | 10 | 0.512 | 0.512 | R | IC | 544-63-8
|  | Palmitic Acid | 5.649 | 32.92 | 25 | 10 | 13.168 | 13.168 | NR | IC | 57-10-3
|  | Oleic Acid | 6.062 | 35.82 | 25 | 10 | 14.328 | 14.328 | NR | IC | 112-80-1
|  | Stearic Acid | 6.822 | 59.17 | 25 | 10 | 23.668 | 23.668 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.651 | 2.89 | 25 | 10 | 1.156 | 1.156 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.108 | 16.01 | 25 | 10 | 6.404 | 6.404 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.366 | 11.84 | 25 | 10 | 4.736 | 4.736 | NR | IC | 31570-04-4
| Accu-Chek S1 | BHT | 4.097 | 0.72 | 25 | 100 | 2.88 | 2.88 | NR | IC | 128-37-0
|  | Palmitic Acid | 5.655 | 4.33 | 25 | 100 | 17.32 | 17.32 | R | IC | 57-10-3
|  | Oleic Acid | 6.062 | 4.31 | 25 | 100 | 17.24 | 17.24 | R | IC | 112-80-1
|  | Stearic Acid | 6.83 | 7.14 | 25 | 100 | 28.56 | 28.56 | R | IC | 57-11-4
|  | Irganox 1010 | 9.653 | 0.16 | 25 | 100 | 0.64 | 0.64 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.109 | 1.59 | 25 | 100 | 6.36 | 6.36 | R | IC | 2082-79-3
|  | Irgafos 168 | 12.37 | 1.02 | 25 | 100 | 4.08 | 4.08 | R | IC | 31570-04-4
| Accu-Chek S2 | BHT | 4.181 | 9.31 | 25 | 10 | 3.724 | 3.724 | R | IC | 128-37-0
|  | Myristic Acid | 4.689 | 1.31 | 25 | 10 | 0.524 | 0.524 | R | IC | 544-63-8
|  | Palmitic Acid | 5.884 | 30.76 | 25 | 10 | 12.304 | 12.304 | NR | IC | 57-10-3
|  | Oleic Acid | 6.306 | 36.76 | 25 | 10 | 14.704 | 14.704 | NR | IC | 112-80-1
|  | Stearic Acid | 7.098 | 59.12 | 25 | 10 | 23.648 | 23.648 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.789 | 2.47 | 25 | 10 | 0.988 | 0.988 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.317 | 15.19 | 25 | 10 | 6.076 | 6.076 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.532 | 10.02 | 25 | 10 | 4.008 | 4.008 | NR | IC | 31570-04-4
| Accu-Chek S2 | BHT | 4.106 | 0.79 | 25 | 100 | 3.16 | 3.16 | NR | IC | 128-37-0
|  | Palmitic Acid | 5.654 | 4.29 | 25 | 100 | 17.16 | 17.16 | R | IC | 57-10-3
|  | Oleic Acid | 6.069 | 4.45 | 25 | 100 | 17.8 | 17.8 | R | IC | 112-80-1
|  | Stearic Acid | 6.873 | 7.27 | 25 | 100 | 29.08 | 29.08 | R | IC | 57-11-4
|  | Irganox 1010 | 9.654 | 0.17 | 25 | 100 | 0.68 | 0.68 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.118 | 1.61 | 25 | 100 | 6.44 | 6.44 | R | IC | 2082-79-3
|  | Irgafos 168 | 12.38 | 1.17 | 25 | 100 | 4.68 | 4.68 | R | IC | 31570-04-4
| Accu-Chek S3 | BHT | 4.224 | 11.62 | 25 | 10 | 4.648 | 4.648 | NR | IC | 128-37-0
|  | Myristic Acid | 4.715 | 1.11 | 25 | 10 | 0.444 | 0.444 | R | IC | 544-63-8
|  | Palmitic Acid | 5.885 | 28.27 | 25 | 10 | 11.308 | 11.308 | NR | IC | 57-10-3
|  | Oleic Acid | 6.306 | 34.3 | 25 | 10 | 13.72 | 13.72 | NR | IC | 112-80-1
|  | Stearic Acid | 7.095 | 48.43 | 25 | 10 | 19.372 | 19.372 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.793 | 3.12 | 25 | 10 | 1.248 | 1.248 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.316 | 14.79 | 25 | 10 | 5.916 | 5.916 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.523 | 9.04 | 25 | 10 | 3.616 | 3.616 | R | IC | 31570-04-4
| Accu-Chek S3 | BHT | 4.106 | 1.16 | 25 | 100 | 4.64 | 4.64 | R | IC | 128-37-0
|  | Palmitic Acid | 5.646 | 3.85 | 25 | 100 | 15.4 | 15.4 | R | IC | 57-10-3
|  | Oleic Acid | 6.061 | 4.34 | 25 | 100 | 17.36 | 17.36 | R | IC | 112-80-1
|  | Stearic Acid | 6.846 | 5.65 | 25 | 100 | 22.6 | 22.6 | R | IC | 57-11-4
|  | Irganox 1010 | 9.667 | 0.25 | 25 | 100 | 1 | 1 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.121 | 1.64 | 25 | 100 | 6.56 | 6.56 | R | IC | 2082-79-3
|  | Irgafos 168 | 12.384 | 1.02 | 25 | 100 | 4.08 | 4.08 | NR | IC | 31570-04-4
| Neria S1 | BHT | 4.104 | 19.87 | 25 | 10 | 7.948 | 7.948 | NR | IC | 128-37-0
|  | Myristic Acid | 4.507 | 1.73 | 25 | 10 | 0.692 | 0.692 | R | IC | 544-63-8
|  | Palmitic Acid | 5.647 | 41.39 | 25 | 10 | 16.556 | 16.556 | NR | IC | 57-10-3
|  | Oleic Acid | 6.052 | 63.86 | 25 | 10 | 25.544 | 25.544 | NR | IC | 112-80-1
|  | Stearic Acid | 6.828 | 76.18 | 25 | 10 | 30.472 | 30.472 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.658 | 4.2 | 25 | 10 | 1.68 | 1.68 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.112 | 16.47 | 25 | 10 | 6.588 | 6.588 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.363 | 21.24 | 25 | 10 | 8.496 | 8.496 | NR | IC | 31570-04-4
| Neria S1 | BHT | 4.107 | 2.34 | 25 | 100 | 9.36 | 9.36 | R | IC | 128-37-0
|  | Palmitic Acid | 5.652 | 7.09 | 25 | 100 | 28.36 | 28.36 | R | IC | 57-10-3
|  | Oleic Acid | 6.059 | 10.12 | 25 | 100 | 40.48 | 40.48 | NR | IC | 112-80-1
|  | Stearic Acid | 6.837 | 11.72 | 25 | 100 | 46.88 | 46.88 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.676 | 0.39 | 25 | 100 | 1.56 | 1.56 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.115 | 2.06 | 25 | 100 | 8.24 | 8.24 | R | IC | 2082-79-3
|  | Irgafos 168 | 12.377 | 1.91 | 25 | 100 | 7.64 | 7.64 | R | IC | 31570-04-4
| Neria S1 | BHT | 4.272 | 0.47 | 25 | 200 | 3.76 | 3.76 | NR | IC | 128-37-0
|  | Palmitic Acid | 6.023 | 2.95 | 25 | 200 | 23.6 | 23.6 | NR | IC | 57-10-3
|  | Oleic Acid | 6.433 | 4.25 | 25 | 200 | 34 | 34 | R | IC | 112-80-1
|  | Stearic Acid | 7.246 | 4.95 | 25 | 200 | 39.6 | 39.6 | R | IC | 57-11-4
|  | Irganox 1010 | 9.87 | 0.1 | 25 | 200 | 0.8 | 0.8 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.443 | 0.87 | 25 | 200 | 6.96 | 6.96 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.634 | 0.95 | 25 | 200 | 7.6 | 7.6 | NR | IC | 31570-04-4
| Neria S2 | BHT | 4.214 | 20.74 | 25 | 10 | 8.296 | 8.296 | NR | IC | 128-37-0
|  | Myristic Acid | 4.717 | 1.62 | 25 | 10 | 0.648 | 0.648 | R | IC | 544-63-8
|  | Palmitic Acid | 5.892 | 41.55 | 25 | 10 | 16.62 | 16.62 | NR | IC | 57-10-3
|  | Oleic Acid | 6.281 | 62.52 | 25 | 10 | 25.008 | 25.008 | NR | IC | 112-80-1
|  | Stearic Acid | 7.107 | 74.43 | 25 | 10 | 29.772 | 29.772 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.797 | 4.01 | 25 | 10 | 1.604 | 1.604 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.319 | 15.94 | 25 | 10 | 6.376 | 6.376 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.541 | 16.81 | 25 | 10 | 6.724 | 6.724 | NR | IC | 31570-04-4
| Neria S2 | BHT | 4.099 | 1.93 | 25 | 100 | 7.72 | 7.72 | R | IC | 128-37-0
|  | Palmitic Acid | 5.667 | 5.18 | 25 | 100 | 20.72 | 20.72 | R | IC | 57-10-3
|  | Oleic Acid | 6.068 | 7.13 | 25 | 100 | 28.52 | 28.52 | R | IC | 112-80-1
|  | Stearic Acid | 6.846 | 8.48 | 25 | 100 | 33.92 | 33.92 | R | IC | 57-11-4
|  | Irganox 1010 | 9.669 | 0.41 | 25 | 100 | 1.64 | 1.64 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.131 | 1.65 | 25 | 100 | 6.6 | 6.6 | R | IC | 2082-79-3
|  | Irgafos 168 | 12.38 | 1.88 | 25 | 100 | 7.52 | 7.52 | R | IC | 31570-04-4
| Neria S3 | BHT | 4.236 | 18.18 | 25 | 10 | 7.272 | 7.272 | NR | IC | 128-37-0
|  | Myristic Acid | 4.723 | 1.4 | 25 | 10 | 0.56 | 0.56 | R | IC | 544-63-8
|  | Palmitic Acid | 5.896 | 35.32 | 25 | 10 | 14.128 | 14.128 | NR | IC | 57-10-3
|  | Oleic Acid | 6.322 | 53.82 | 25 | 10 | 21.528 | 21.528 | NR | IC | 112-80-1
|  | Stearic Acid | 7.107 | 64.36 | 25 | 10 | 25.744 | 25.744 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.802 | 3.2 | 25 | 10 | 1.28 | 1.28 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.334 | 12.35 | 25 | 10 | 4.94 | 4.94 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.54 | 13.4 | 25 | 10 | 5.36 | 5.36 | NR | IC | 31570-04-4
| Neria S3 | BHT | 4.117 | 1.92 | 25 | 100 | 7.68 | 7.68 | R | IC | 128-37-0
|  | Palmitic Acid | 5.672 | 4.12 | 25 | 100 | 16.48 | 16.48 | R | IC | 57-10-3
|  | Oleic Acid | 6.076 | 5.93 | 25 | 100 | 23.72 | 23.72 | R | IC | 112-80-1
|  | Stearic Acid | 6.846 | 7.45 | 25 | 100 | 29.8 | 29.8 | R | IC | 57-11-4
|  | Irganox 1010 | 9.682 | 0.22 | 25 | 100 | 0.88 | 0.88 | NR | IC | 98584-37-3
|  | Irganox 1076 | 11.134 | 1.27 | 25 | 100 | 5.08 | 5.08 | R | IC | 2082-79-3
|  | Irgafos 168 | 12.39 | 1.4 | 25 | 100 | 5.6 | 5.6 | R | IC | 31570-04-4
| Syringe S1 | BHT | 4.106 | 82.1 | 25 | 10 | 32.84 | 32.84 | NR | IC | 128-37-0
|  | Myristic Acid | 4.547 | 3.38 | 25 | 10 | 1.352 | 1.352 | R | IC | 544-63-8
|  | A.O. 2246 | 5.119 | 0.8 | 25 | 10 | 0.32 | 0.32 | R | IC | 119-47-1
|  | Palmitic Acid | 5.66 | 30.53 | 25 | 10 | 12.212 | 12.212 | NR | IC | 57-10-3
|  | Oleic Acid | 6.014 | 60.07 | 25 | 10 | 24.028 | 24.028 | NR | IC | 112-80-1
|  | Stearic Acid | 6.829 | 55.73 | 25 | 10 | 22.292 | 22.292 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.646 | 2.35 | 25 | 10 | 0.94 | 0.94 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.1 | 0.19 | 25 | 10 | 0.076 | 0.076 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.364 | 10.25 | 25 | 10 | 4.1 | 4.1 | NR | IC | 31570-04-4
| Syringe S1 | BHT | 4.085 | 11.09 | 25 | 100 | 44.36 | 44.36 | NR | IC | 128-37-0
|  | A.O. 2246 | 5.071 | 0.038 | 25 | 100 | 0.152 | 0.152 | NR | IC | 119-47-1
|  | Palmitic Acid | 5.606 | 4.22 | 25 | 100 | 16.88 | 16.88 | R | IC | 57-10-3
|  | Oleic Acid | 5.962 | 9.44 | 25 | 100 | 37.76 | 37.76 | R | IC | 112-80-1
|  | Stearic Acid | 6.779 | 10.99 | 25 | 100 | 43.96 | 43.96 | R | IC | 57-11-4
|  | Irganox 1010 | 9.626 | 0.19 | 25 | 100 | 0.76 | 0.76 | NR | IC | 98584-37-3
|  | Irgafos 168 | 12.332 | 1.4 | 25 | 100 | 5.6 | 5.6 | R | IC | 31570-04-4
| Syringe S1 | BHT | 4.279 | 5.21 | 25 | 200 | 41.68 | 41.68 | R | IC | 128-37-0
|  | Palmitic Acid | 6.012 | 2.12 | 25 | 200 | 16.96 | 16.96 | NR | IC | 57-10-3
|  | Oleic Acid | 6.385 | 4.83 | 25 | 200 | 38.64 | 38.64 | NR | IC | 112-80-1
|  | Stearic Acid | 7.246 | 5.92 | 25 | 200 | 47.36 | 47.36 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.871 | 0.03 | 25 | 200 | 0.24 | 0.24 | NR | IC | 98584-37-3
|  | Irgafos 168 | 12.622 | 0.69 | 25 | 200 | 5.52 | 5.52 | NR | IC | 31570-04-4
| Syringe S2 | BHT | 4.1 | 121.26 | 25 | 10 | 48.504 | 48.504 | NR | IC | 128-37-0
|  | Myristic Acid | 4.513 | 4.82 | 25 | 10 | 1.928 | 1.928 | R | IC | 544-63-8
|  | A.O. 2246 | 5.096 | 1.28 | 25 | 10 | 0.512 | 0.512 | R | IC | 119-47-1
|  | Palmitic Acid | 5.632 | 41.28 | 25 | 10 | 16.512 | 16.512 | NR | IC | 57-10-3
|  | Oleic Acid | 6.001 | 88.92 | 25 | 10 | 35.568 | 35.568 | NR | IC | 112-80-1
|  | Stearic Acid | 6.803 | 67.94 | 25 | 10 | 27.176 | 27.176 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.645 | 3.39 | 25 | 10 | 1.356 | 1.356 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.105 | 0.21 | 25 | 10 | 0.084 | 0.084 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.364 | 16.57 | 25 | 10 | 6.628 | 6.628 | NR | IC | 31570-04-4
| Syringe S2 | BHT | 4.109 | 13.13 | 25 | 100 | 52.52 | 52.52 | NR | IC | 128-37-0
|  | Myristic Acid | 4.506 | 0.57 | 25 | 100 | 2.28 | 2.28 | NR | IC | 544-63-8
|  | A.O. 2246 | 5.108 | 0.05 | 25 | 100 | 0.2 | 0.2 | NR | IC | 119-47-1
|  | Palmitic Acid | 5.65 | 5.17 | 25 | 100 | 20.68 | 20.68 | R | IC | 57-10-3
|  | Oleic Acid | 6.001 | 10.81 | 25 | 100 | 43.24 | 43.24 | NR | IC | 112-80-1
|  | Stearic Acid | 6.831 | 13.3 | 25 | 100 | 53.2 | 53.2 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.659 | 0.31 | 25 | 100 | 1.24 | 1.24 | NR | IC | 98584-37-3
|  | Irgafos 168 | 12.384 | 1.63 | 25 | 100 | 6.52 | 6.52 | R | IC | 31570-04-4
| Syringe S2 | BHT | 4.26 | 8.01 | 25 | 200 | 64.08 | 64.08 | R | IC | 128-37-0
|  | Palmitic Acid | 5.962 | 3.15 | 25 | 200 | 25.2 | 25.2 | NR | IC | 57-10-3
|  | Oleic Acid | 6.327 | 6.08 | 25 | 200 | 48.64 | 48.64 | R | IC | 112-80-1
|  | Stearic Acid | 7.181 | 7.71 | 25 | 200 | 61.68 | 61.68 | R | IC | 57-11-4
|  | Irganox 1010 | 9.838 | 0.06 | 25 | 200 | 0.48 | 0.48 | NR | IC | 98584-37-3
|  | Irgafos 168 | 12.581 | 0.96 | 25 | 200 | 7.68 | 7.68 | NR | IC | 31570-04-4
| Syringe S3 | BHT | 4.097 | 97.55 | 25 | 10 | 39.02 | 39.02 | NR | IC | 128-37-0
|  | Myristic Acid | 4.51 | 4.1 | 25 | 10 | 1.64 | 1.64 | R | IC | 544-63-8
|  | A.O. 2246 | 5.095 | 1.21 | 25 | 10 | 0.484 | 0.484 | R | IC | 119-47-1
|  | Palmitic Acid | 5.639 | 32.5 | 25 | 10 | 13 | 13 | NR | IC | 57-10-3
|  | Oleic Acid | 5.991 | 70.6 | 25 | 10 | 28.24 | 28.24 | NR | IC | 112-80-1
|  | Stearic Acid | 6.814 | 56.56 | 25 | 10 | 22.624 | 22.624 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.648 | 2.83 | 25 | 10 | 1.132 | 1.132 | R | IC | 98584-37-3
|  | Irganox 1076 | 11.101 | 0.19 | 25 | 10 | 0.076 | 0.076 | NR | IC | 2082-79-3
|  | Irgafos 168 | 12.36 | 13.35 | 25 | 10 | 5.34 | 5.34 | NR | IC | 31570-04-4
| Syringe S3 | BHT | 4.109 | 12.44 | 25 | 100 | 49.76 | 49.76 | NR | IC | 128-37-0
|  | A.O. 2246 | 5.116 | 0.07 | 25 | 100 | 0.28 | 0.28 | NR | IC | 119-47-1
|  | Palmitic Acid | 5.676 | 5.07 | 25 | 100 | 20.28 | 20.28 | R | IC | 57-10-3
|  | Oleic Acid | 6.031 | 11.13 | 25 | 100 | 44.52 | 44.52 | NR | IC | 112-80-1
|  | Stearic Acid | 6.851 | 13.42 | 25 | 100 | 53.68 | 53.68 | NR | IC | 57-11-4
|  | Irganox 1010 | 9.67 | 0.27 | 25 | 100 | 1.08 | 1.08 | NR | IC | 98584-37-3
|  | Irgafos 168 | 12.392 | 1.88 | 25 | 100 | 7.52 | 7.52 | R | IC | 31570-04-4
| Syringe S3 | BHT | 4.29 | 4.21 | 25 | 200 | 33.68 | 33.68 | R | IC | 128-37-0
|  | Palmitic Acid | 5.963 | 1.88 | 25 | 200 | 15.04 | 15.04 | NR | IC | 57-10-3
|  | Oleic Acid | 6.328 | 3.733 | 25 | 200 | 29.864 | 29.864 | R | IC | 112-80-1
|  | Stearic Acid | 7.173 | 4.42 | 25 | 200 | 35.36 | 35.36 | R | IC | 57-11-4
|  | Irganox 1010 | 9.829 | 0.05 | 25 | 200 | 0.4 | 0.4 | NR | IC | 98584-37-3
|  | Irgafos 168 | 12.58 | 0.75 | 25 | 200 | 6 | 6 | NR | IC | 31570-04-4
BRL - Below reporting Limit (<1 ug/mL)
Compounds observed in blank not reported, as they are not test article/control related.
Corrected Resulted (ug/mL) = (Instrument Result (ug/mL) * Dilution Factor)/Concentration Factor
Blank/Control Corrected Result (ug/mL) = Correct Result (ug/mL) - Corrected Result Corrected for Blank/Control (ug/mL)
Concentration Factor = Initial Extraction Volume/Concentrated Volume
Highlighted cells are peaks that were integrated but not automatically reported. Need to be analyst verified.

## LCUV
15-03379-N1 |  | LC/UV
Peak
|  | Additional Peaks | Ret Time | Wavelenth(s) (nm) | Area | Results (ug/mL) | Area of Standard
Drug | Control | N/A | N/A | N/A | N/A | N/A | N/A
| Accu-Check d5 S1 | None | N/A | N/A | N/A | N/A | N/A
| Accu-Check d5 S2 | None | N/A | N/A | N/A | N/A | N/A
| Accu-Check d5 S3 | None | N/A | N/A | N/A | N/A | N/A
| Neria d5 S1 | Yes | 7.522 | 205 | 1.1917 | 0.011965 | 995.9467
| Neria d5 S2 | Yes | 7.521 | 205 | 16.9928 | 0.17402 | 976.4641
| Neria d5 S3 | Yes | 2.081 | 215270320 | 1846.7878 | 18.5648 | 994.7817
|  |  | 7.518 | 205 | 257.0289 | 2.5838
| Syringe d5 S1 | None | N/A | N/A | N/A | N/A | N/A
| Syringe d5 S2 | None | N/A | N/A | N/A | N/A | N/A
| Syringe d5 S3 | None | N/A | N/A | N/A | N/A | N/A
15-03379-N1 |  | LC/UV
Peak
|  | Additional Peaks | Ret Time | Wavelenth(s) (nm) | Area | Results (ug/mL) | Area of Standard
Acidified Water | Control | N/A | N/A | N/A | N/A | N/A | N/A
| Accu-Check S1 | None | N/A | N/A | N/A | N/A | N/A
| Accu-Check S2 | None | N/A | N/A | N/A | N/A | N/A
| Accu-Check S3 | None | N/A | N/A | N/A | N/A | N/A
| Neria S1 | None | N/A | N/A | N/A | N/A | N/A
| Neria S2 | None | N/A | N/A | N/A | N/A | N/A
| Neria S3 | None | N/A | N/A | N/A | N/A | N/A
| Syringe S1 | Yes | 2.297 | 205275 | 395.3709 | 3.6694 | 1077.4829
|  |  | 3.972 | 205 | 147.3539 | 1.3676
| Syringe S2 | Yes | 2.283 | 205275 | 290.5882 | 3.0951 | 938.858
|  |  | 3.976 | 205 | 109.6379 | 1.1678
| Syringe S3 | Yes | 2.23 | 205275 | 286.3125 | 3.2291 | 886.6661
15-03379-N1
Peak
|  | Additional Peaks | Ret Time | Wavelenth(s) (nm) | Area | Results (ug/mL) | Area of Standard
IPA | Control | N/A | N/A | N/A | N/A | N/A | N/A
| Accu-Check d10 S1 | Yes | 1.475 | 205250 | 5786.3198 | 61.1497 | 946.2548
|  |  | 1.674 | 205250 | 5163.9448 | 54.5725
|  |  | 1.837 | 205250 | 10146.8 | 107.2312
|  |  | 1.954 | 205245 | 10664.1 | 112.698
|  |  | 2.049 | 210250 | 2342.4353 | 24.7548
|  |  | 2.349 | 210250 | 2888.3291 | 30.5238
|  |  | 2.526 | 210250 | 5361.0366 | 56.6553
|  |  | 2.852 | 210250305 | 152.5183 | 1.6118
|  |  | 3.171 | 210250 | 298.5899 | 3.1555
|  |  | 3.57 | 210250 | 557.1083 | 5.8875
|  |  | 4.647 | 205 | 157.022 | 1.6594
|  |  | 5.011 | 210245 | 438.5385 | 4.6345
|  |  | 5.414 | 210 | 2119.3782 | 22.3975
|  |  | 6.04 | 205230 | 160.1081 | 1.692
|  |  | 7.375 | 210 | 354.216 | 3.7433
|  |  | 10.28 | 205 | 333.0251 | 3.5194
|  |  | 10.525 | 205 | 245.0538 | 2.5897
Results (ug/mL) = (Area Peak*(Conc standard (ug/mL)))/Area Standard
Conc of Standard = 10 ug/mL
15-03379-N1 |  | LC/UV
Peak
|  | Additional Peaks | Ret Time | Wavelenth(s) (nm) | Area | Results (ug/mL) | Area of Standard
IPA | Control | N/A | N/A | N/A | N/A | N/A | N/A
(cont)
| Accu-Check d10 S2 | Yes | 1.485 | 205250 | 6278.9507 | 64.4554 | 974.1549
|  |  | 1.66 | 205250 | 4691.75 | 48.1623
|  |  | 1.8 | 205250 | 4559.1548 | 46.8011
|  |  | 1.846 | 205250 | 5344.0181 | 54.858
|  |  | 1.98 | 205250 | 11362.4 | 116.6385
|  |  | 2.14 | 210250 | 5078.6055 | 52.1334
|  |  | 2.365 | 210250 | 3629.866 | 37.2617
|  |  | 2.559 | 210250 | 6102.4409 | 62.6434
|  |  | 2.905 | 210250305 | 484.2056 | 4.9705
|  |  | 3.233 | 210250 | 290.9395 | 2.9866
|  |  | 3.658 | 210250 | 608.0836 | 6.2422
|  |  | 4.781 | 205 | 177.8525 | 1.8257
|  |  | 5.145 | 210245 | 446.5077 | 4.5835
|  |  | 5.605 | 210 | 1598.293 | 16.407
|  |  | 5.797 | 210235285 | 506.9203 | 5.2037
|  |  | 6.273 | 205230 | 153.8695 | 1.5795
|  |  | 7.623 | 210 | 381.0237 | 3.9113
|  |  | 10.52 | 205235 | 350.0215 | 3.5931
|  |  | 10.705 | 205 | 253.3386 | 2.6006
| Accu-Check d10 S3 | Yes | 1.479 | 205250 | 5559.1646 | 57.4417 | 967.7933
|  |  | 1.71 | 205250 | 4626.5478 | 47.8051
|  |  | 1.783 | 205250 | 2976.4482 | 30.755
|  |  | 1.856 | 205250 | 6483.6846 | 66.9945
|  |  | 1.997 | 205250 | 12238.84 | 126.4613
|  |  | 2.154 | 210250 | 5153.0161 | 53.245
|  |  | 2.376 | 210250 | 3271.5002 | 33.8037
|  |  | 2.574 | 210250 | 5467.1992 | 56.4914
|  |  | 2.922 | 210250305 | 354.2534 | 3.6604
|  |  | 3.253 | 210250 | 285.7311 | 2.9524
|  |  | 3.674 | 210250 | 663.3388 | 6.8541
|  |  | 4.783 | 205 | 151.331 | 1.5637
|  |  | 5.151 | 210245 | 455.7409 | 4.7091
|  |  | 5.606 | 210 | 1494.4021 | 15.4413
|  |  | 5.796 | 210235285 | 505.3993 | 5.2222
|  |  | 6.266 | 205230 | 149.8899 | 1.5488
|  |  | 7.621 | 210 | 378.1707 | 3.9076
|  |  | 10.522 | 205235 | 342.4802 | 3.5388
|  |  | 10.705 | 205 | 261.5984 | 2.703
| Neria d10 S1 |  | 1.488 | 210250 | 8067.3804 | 83.6943 | 963.9108
|  |  | 1.667 | 210250 | 6345.169 | 65.8273
|  |  | 1.842 | 210250 | 13915 | 144.3598
|  |  | 1.958 | 215250 | 22658.7 | 235.0705
|  |  | 2.345 | 210250 | 2794.864 | 28.995
|  |  | 2.524 | 205250 | 5794.0625 | 60.1099
|  |  | 2.839 | 205250305 | 156.2603 | 1.6211
|  |  | 3.171 | 210250 | 111.7943 | 1.1598
|  |  | 3.569 | 205245 | 539.9351 | 5.6015
|  |  | 5.009 | 210245 | 445.2487 | 4.6192
|  |  | 5.411 | 205 | 2283.6697 | 23.6917
|  |  | 6.028 | 210230 | 274.4931 | 2.8477
|  |  | 7.146 | 210 | 150.2551 | 1.5588
|  |  | 7.376 | 205 | 357.0666 | 3.7044
|  |  | 8.893 | 205 | 115.9113 | 1.2025
|  |  | 10.286 | 205 | 440.9551 | 4.5746
|  |  | 10.527 | 205 | 556.779 | 5.7763
Results (ug/mL) = (Area Peak*(Conc standard (ug/mL)))/Area Standard
Conc of Standard = 10 ug/mL
15-03379-N1 |  | LC/UV
IPA |  |  |  |  | Peak
(cont) |  | Additional Peaks | Ret Time | Wavelenth(s) (nm) | Area | Results (ug/mL) | Area of Standard
| Neria d10 S2 |  | 1.483 | 210250 | 6412.2593 | 65.4543 | 979.655
|  |  | 1.649 | 210250 | 4504.8901 | 45.9845
|  |  | 1.856 | 210250 | 12692.5 | 129.5609
|  |  | 1.997 | 215250 | 21759.7 | 222.116
|  |  | 2.372 | 210250 | 4181.8442 | 42.6869
|  |  | 2.572 | 205250 | 6447.3984 | 65.813
|  |  | 2.929 | 205250305 | 855.4208 | 8.7319
|  |  | 3.253 | 210250 | 143.7336 | 1.4672
|  |  | 3.677 | 205245 | 505.4863 | 5.1598
|  |  | 5.158 | 210245 | 414.0732 | 4.2267
|  |  | 5.614 | 205 | 1629.1365 | 16.6297
|  |  | 5.807 | 205235280 | 555.5878 | 5.6713
|  |  | 6.276 | 210230 | 232.511 | 2.3734
|  |  | 7.399 | 210 | 124.532 | 1.2712
|  |  | 7.637 | 205 | 272.1158 | 2.7777
|  |  | 9.163 | 205 | 110.1046 | 1.1239
|  |  | 10.529 | 205 | 355.2087 | 3.6259
|  |  | 10.712 | 205 | 587.1279 | 5.9932
| Neria d10 S3 |  | 1.469 | 210250 | 5525.7388 | 55.9215 | 988.1243
|  |  | 1.713 | 210250 | 4756.8828 | 48.1405
|  |  | 1.857 | 210250 | 10622.8 | 107.5047
|  |  | 1.997 | 215250 | 17001.5 | 172.0583
|  |  | 2.371 | 210250 | 3153.8552 | 31.9176
|  |  | 2.573 | 205250 | 4687.6338 | 47.4397
|  |  | 2.93 | 205250305 | 303.4058 | 3.0705
|  |  | 3.253 | 210250 | 119.5575 | 1.2099
|  |  | 3.676 | 205245 | 436.2992 | 4.4154
|  |  | 5.16 | 210245 | 327.8251 | 3.3177
|  |  | 5.615 | 205 | 1319.6763 | 13.3554
|  |  | 5.807 | 205235280 | 406.7545 | 4.1164
|  |  | 6.271 | 210230 | 185.6983 | 1.8793
|  |  | 7.388 | 210 | 102.8043 | 1.0404
|  |  | 7.624 | 205 | 248.7387 | 2.5173
|  |  | 9.159 | 205 | 100.4052 | 1.0161
|  |  | 10.535 | 205 | 335.1074 | 3.3913
|  |  | 10.716 | 205 | 456.3095 | 4.6179
| Syringe d10 S1 |  | 1.292 | 215 | 361.545 | 2.9565 | 1222.8782
|  |  | 1.459 | 205 | 695.1936 | 5.6849
|  |  | 2 | 220 | 827.8578 | 6.7697
|  |  | 2.324 | 210275 | 680.0925 | 5.5614
|  |  | 2.586 | 210230275 | 312.8811 | 2.5586
|  |  | 2.839 | 205230280 | 205.0135 | 1.6765
|  |  | 3.192 | 210280335 | 1981.3962 | 16.2027
|  |  | 3.404 | 210260 | 3303.1877 | 27.0116
|  |  | 4.485 | 210285 | 4439.5674 | 36.3042
|  |  | 5.031 | 205280 | 658.8431 | 5.3876
|  |  | 5.188 | 205 | 604.2109 | 4.9409
|  |  | 5.71 | 205280 | 853.1303 | 6.9764
|  |  | 6.092 | 205250 | 1915.8899 | 15.6671
|  |  | 6.564 | 205285 | 3303.1223 | 27.011
|  |  | 6.803 | 210285 | 9914.2519 | 81.0731
|  |  | 7.137 | 210280 | 678.6424 | 5.5496
|  |  | 7.567 | 205, 250,330,350 | 779.8918 | 6.3775
|  |  | 8.121 | 205225280 | 665.7694 | 5.4443
|  |  | 8.349 | 205280 | 608.43 | 4.9754
|  |  | 8.588 | 205280 | 362.3 | 2.9627
|  |  | 8.813 | 205280 | 775.4938 | 6.3415
|  |  | 8.932 | 205285 | 1453.1245 | 11.8828
|  |  | 9.134 | 205285 | 683.7531 | 5.5913
|  |  | 9.372 | 205280 | 257.0892 | 2.1023
|  |  | 9.78 | 205225285 | 225.9168 | 1.8474
|  |  | 10.254 | 205225285 | 1120.9156 | 9.1662
|  |  | 10.519 | 210 | 2113.6272 | 17.284
|  |  | 11.51 | 210230 | 174.9359 | 1.4305
Results (ug/mL) = (Area Peak*(Conc standard (ug/mL)))/Area Standard
Conc of Standard = 10 ug/mL
15-03379-N1 |  | LC/UV
IPA |  |  |  |  | Peak
(cont) |  | Additional Peaks | Ret Time | Wavelenth(s) (nm) | Area | Results (ug/mL) | Area of Standard
| Syringe d10 S2 |  | 1.3 | 215 | 461.3335 | 2.7111 | 1701.6414
|  |  | 1.451 | 205 | 670.1713 | 3.9384
|  |  | 2.081 | 220 | 1364.1337 | 8.0166
|  |  | 2.316 | 210275 | 1206.0078 | 7.0873
|  |  | 2.578 | 210230275 | 762.4969 | 4.4809
|  |  | 2.831 | 205230280 | 516.6003 | 3.0359
|  |  | 3.176 | 210280335 | 2948.3826 | 17.3267
|  |  | 3.389 | 210260 | 4509.3682 | 26.5001
|  |  | 4.468 | 210285 | 5973.7441 | 35.1058
|  |  | 5.008 | 205280 | 907.8638 | 5.3352
|  |  | 5.17 | 205 | 784.929 | 4.6128
|  |  | 5.693 | 205280 | 1071.8004 | 6.2986
|  |  | 6.074 | 205250 | 2425.9739 | 14.2567
|  |  | 6.549 | 205285 | 4356.8657 | 25.6039
|  |  | 6.787 | 210285 | 12573.8 | 73.8922
|  |  | 7.127 | 210280 | 1096.6846 | 6.4449
|  |  | 7.555 | 205, 250,330,350 | 774.8123 | 4.5533
|  |  | 8.111 | 205225280 | 924.2072 | 5.4313
|  |  | 8.34 | 205280 | 584.6561 | 3.4358
|  |  | 8.584 | 205280 | 697.5693 | 4.0994
|  |  | 8.811 | 205280 | 1099.979 | 6.4642
|  |  | 8.93 | 205285 | 2083.4724 | 12.2439
|  |  | 9.132 | 205285 | 1022.1406 | 6.0068
|  |  | 9.365 | 205280 | 218.9386 | 1.2866
|  |  | 9.78 | 205225280 | 261.4516 | 1.5365
|  |  | 10.256 | 205225285 | 1559.2766 | 9.1634
|  |  | 10.52 | 210 | 3054.4265 | 17.9499
|  |  | 11.516 | 205230 | 296.4735 | 1.7423
| Syringe d10 S3 |  | 1.299 | 215 | 400.7784 | 2.7114 | 1478.126
|  |  | 1.467 | 205 | 687.1034 | 4.6485
|  |  | 2.085 | 220 | 1183.381 | 8.006
|  |  | 2.314 | 210275 | 1081.6969 | 7.318
|  |  | 2.577 | 210230275 | 696.8646 | 4.7145
|  |  | 2.831 | 205230280 | 479.7308 | 3.2455
|  |  | 3.175 | 210280335 | 2560.6228 | 17.3234
|  |  | 3.388 | 210260 | 3912.3136 | 26.4681
|  |  | 4.042 | 210285 | 6532.5745 | 44.195
|  |  | 4.468 | 210285 | 5097.6001 | 34.4869
|  |  | 5.012 | 205280 | 761.5433 | 5.1521
|  |  | 5.172 | 205 | 671.1638 | 4.5406
|  |  | 5.698 | 205280 | 944.5768 | 6.3904
|  |  | 6.08 | 205250 | 2129.5369 | 14.407
|  |  | 6.555 | 205285 | 3824.1062 | 25.8713
|  |  | 6.792 | 210285 | 11272.2 | 76.2601
|  |  | 7.129 | 210280 | 852.8559 | 5.7698
|  |  | 7.555 | 205, 250,330,350 | 787.301 | 5.3263
|  |  | 8.108 | 205225280 | 756.8372 | 5.1202
|  |  | 8.338 | 205280 | 555.621 | 3.759
|  |  | 8.582 | 205280 | 482.3854 | 3.2635
|  |  | 8.811 | 205280 | 958.7375 | 6.4862
|  |  | 8.928 | 205285 | 1749.1986 | 11.8339
|  |  | 9.133 | 205285 | 823.2469 | 5.5695
|  |  | 9.363 | 205280 | 177.9659 | 1.204
|  |  | 9.78 | 205225280 | 242.7226 | 1.6421
|  |  | 10.259 | 205225285 | 1313.4944 | 8.8862
|  |  | 10.522 | 210 | 2482.9458 | 16.7979
|  |  | 11.518 | 205230 | 262.9536 | 1.779
Results (ug/mL) = (Area Peak*(Conc standard (ug/mL)))/Area Standard
Conc of Standard = 10 ug/mL

## pH - Conductivity
15-03379-N1 |  | pH
15 minute
|  | Result | Average | St dev
WFI | Control | 6.98 | 6.98 | N/A
| Sample 1 | 6.54 | 6.5833 | 0.29738
| Sample 2 | 6.31
| Sample 3 | 6.9
15-03379-N1 |  | Conductivity
15 minute |  |  | Average
|  | Result (uS/cm) | Result (uS/cm) | St dev
WFI | Control | 1.93 | 1.93 | N/A
| Sample 1 | 2.03 | 2.44 | 0.43715
| Sample 2 | 2.9
| Sample 3 | 2.39

## TOC
15-03379-N1 |  | TOC
15 Minute |  | Instrument | Blank Correct
|  | Result (mg/L) | Result (mg/L) | Average (mg/L) | Std Dev
| Control | 0.2077 | 0.2077 | 0.2077 | N/A
| Sample 1 | 0.3719 | 0.1642 | 0.2602 | 0.083173
| Sample 2 | 0.5135 | 0.3058
| Sample 3 | 0.5183 | 0.3106
Correct Result (ug/mL) = Instrument Result (mg/L) - Control Result (mg/L)

## CAS #
VOC |  | SVOC |  | NVOC
Compound | CAS #
Acetone | 67-64-1 |  |  | Compound | CAS #
Caprolactam | 105-60-2 | Compound | CAS # | Irganox 245 | 36443-68-2
Tert butyl alcohol | 75-65-0 | Diethylphthalate | 84-66-2 | BHT | 128-37-0
|  | Dimethylphthalate | 131-11-3 | Myristic acid | 544-63-8
2 methyl 1 propene | 115-11-7 | bis (2-ethylhexyl) phthalate | 117-81-7 | Oleamide | 301-02-0
ethyl alcohol | 64-17-5 | Bisphenol A | 80-05-7 | A.O. 2246 | 119-47-1
2-ethyl 1 hexanol | 104-76-7 |  |  | Palmitic acid | 57-10-3
Butanal | 123-72-8 | Trioctyl Trimellitate (TOTM) | 3319-31-1 | Oleic acid | 112-80-1
Pentanal | 110-62-3 |  |  | Stearic acid | 57-11-4
Hexanal | 66-25-1 |  |  | Erucamide | 112-84-5
|  |  |  | Irganox 1010 | 98584-37-3
Benzyl alcohol | 100-51-6 |  |  | Irganox 1330 | 1709-70-5
2-butanone | 78-93-3 |  |  | Irganox 1076 | 2082-79-3
2-hexanone | 591-78-6 |  |  | Irgafos 168 | 31570-04-4
|  |  |  | Irgafos 126 | 26741-53-7
|  |  |  | Irganox 3114 | 27676-62-6
1-propene, 2methyl- | 115-11-7
methylene chloride | 75-09-2
acetonitrile | 75-05-8
chloromethane | 74-87-3
bromomethane | 74-83-9
IPA | 67-63-0
tetrahydrofuran | 109-99-9
acrylonitrile | 1424-48-2
styrene | 100-42-5
chloro benzene | 108-90-7
naphthalene | 91-20-3
4-methyl-2-pentanone | 108-10-1
ethylbenzene | 100-41-4
m,p-xylene | 1330-20-7
n-hexane | 110-54-3
toluene | 108-88-3
Benzene | 71-43-2
1,1,1-trichloroethane | 71-55-6
Carbon tetrachloride | 56-23-5
ethyl tert-butyl ether | 637-92-3
1,3,5-trimethylbenzene | 108-67-8
1,2,4-trimethylbenzene | 95-63-6
Chloroform | 67-66-3
2,2,4,6,6-pentamethylheptane | 13475-82-6